## Table 1
LIS
Vendor Interface Document
BECTON, DICKINSON AND COMPANY
7 Loveton Circle Sparks, Maryland 21152 USA
800–638–8663
October 2013
Document Number: L–005933(Q)(17)
Change History
Revision Date Pages Reason
(A) | 10/99 | All | New manual release
(B) | 3/00 | All | New fields added to EpiCenter interface to support Phoenix ID/AST instrument data. Added Isolate results reporting for EpiCenter. Added additional results for tests. And changed Patient and Specimen user defined fields 4 & 5 to free text fields.
(C) | 3/01 | All | Added BD Phoenix 100 section. Updated list of antibiotic and organism codes. EpiCenter uploads chartable rules and special messages. EpiCenter allows blanking of a field. EpiCenter adds fields for excluding isolates from statistics and test complete DT.
(D) | 5/01 | 2, 17, 18, 22, 28, 61-63, 65 | Updated orders record to incorporate critical panel specification. Updated result record to include resistance markers 4 – 10. Changed ‘resistance mechanism’ to ‘resistance marker’ throughout. Changed “BACTEC9000” to “BT9000” in instrument type result field. Fix Phoenix resistance marker result upload example.
(E) | 10/01 | All | Formatting changes
(F) | 3/02 | All | Various updates throughout the document to cleanup references/definitions and lists of information.
(G) | 12/03 | All | Updates for EpiCenter Version 4 Features
(H) | 12/03 | 18,31,63,65-76 | Updates for Phoenix Version 4
(I) | - | - | Skipped
(J) | 2004-04-14 | All | Phoenix Version 4.01Z and EpiCenter 4.10A updates, BD ProbeTec ET and Viper SP updates for new configuration options, status, and assay type list.
(K) | 2005-03-14 | All | Various updates throughout the document to cleanup examples.
(L) | 2008-04-28 | 13, 16, 22, 25, 28, 31, 32, 38, 39, 41, 42 | Updates for FX , Emulating BACTEC 9000, addition of ID test source field, addition of isolate upload examples
(M) | 2008-09-11 | 83--98 | Added section 14.4 describing BACTEC FX standalone
(N) | 2009-02-25 | 22, 30, 66-69 | Added new Viper XTR assay codes for HSV1 and HSV2
(O) | 7/02/2010 06/06/2011 | 25, 27, 37 23, 30, 33-35, 41,47, 101-109 | Add additional value to Time Qualifier field, add note to Admission Date/Time,Collection Date/Time and Received Date/Time fields Updates for Innova instrument
(P) | 2012-07-27 | 72 | Added new Viper XTR assay codes for HPV and TVQx
(Q)(17) | 2013-10-24 | 69-75 | Added additional information for Viper LT
BD is a registered trademark of Becton, Dickinson and Company.
© Copyright Becton, Dickinson and Company, 2004. All rights reserved. No part of this publication may be reproduced, transmitted, transcribed, stored in retrieval systems, or translated into any language or computer language, in any form or by any means, electronic, mechanical, magnetic, optical, chemical, manual, or otherwise, without the prior written permission of BD, 7 Loveton Circle, Sparks, Maryland, 21152, United States of America.
Table of Contents
1. Purpose 7
2. Scope 7
3. Definitions 7
4. Overview 7
4.1. Physical Cable Interface 7
4.2. Communication Protocols 8
4.3. System Configurations 8
4.4. EpiCenter- LIS Communication 10
5. Workflow 10
5.1. Uploaded Test Results 10
5.2. Uploaded Isolate Results 11
5.3. Unsolicited Downloads 14
5.4. File Import 15
5.5. Queries 15
5.6. Error Reporting 15
6. Protocol Specifications 16
6.1. Physical Protocols 16
6.1.1. Memory limitations 16
6.1.2. Unpacked Frames 16
6.1.3. BDMODEM 16
6.1.3.1. Emulate BACTEC 9000 LIS Interface 17
6.2. Logical Protocols 17
6.2.1. Unprocessed Records 17
6.2.2. Download Only Fields 18
6.2.3. Communication Errors 18
6.2.4. LIS Codes 18
6.2.5. Blanking fields 19
7. Configurable Options 19
7.1. Port parameters 19
7.2. Auto-Populating Lists 19
7.3. Logical Protocol Parameters 20
7.4. ASTM Field Mapping 20
7.5. User Defined Fields. 20
8. Message Content 21
8.1. Field List 21
8.2. Field Descriptions 24
8.2.1. Header Fields 24
8.2.2. Patient Fields 24
8.2.3. Order Fields 26
8.2.4. Result Fields 28
8.2.5. Comment Fields 34
8.2.6. Request Fields 35
8.2.7. Terminator Fields 37
9. Simulator 38
10. Appendix A - Sample EpiCenter Messages 39
10.1. LIS Order 39
10.2. Host Query 39
10.3. LIS Requests Results 39
10.4. Results 40
10.4.1. BACTEC G&D Test results 40
10.4.2. MGIT 960 G&D Test results: 40
10.4.3. MGIT AST Test Results 40
10.4.4. Innova Results 40
10.4.5. Isolate Results 41
10.4.5.1. Isolate with Expert system changes - Blank MIC, modified SIR 41
10.4.5.2. Isolate with Resistance Markers 42
10.4.5.3. Isolate with Expert Comments 42
10.4.5.4. Isolate with inferred results 43
10.4.5.5. Isolate with multiple test sources 43
10.4.5.6. Isolate with multiple instruments 44
11. Appendix B – Tests and Results 45
12. Appendix C – Antibiotic Abbreviations 48
13. Appendix D – Organism Abbreviations 50
14. Appendix E – BD Instrument Specifics 59
14.1. BACTEC MGIT 960 59
14.1.1. Configurable Options 59
14.1.2. Message Content 59
Field List 59
Field Descriptions 61
Header Fields 61
Patient Fields 61
Order Fields 61
Result Fields 61
Request Fields 62
Terminator Fields 63
14.1.3. Example ASTM 1394 Protocol 64
14.2. BDProbeTec ET, Viper XTR, & Viper LT 67
14.2.1. Configurable Options 67
14.2.2. Message Content 67
Field List 67
Field Descriptions 68
Header Fields 68
Patient Fields 69
Order Fields 69
Result Fields 69
Request Fields 71
Terminator Field 71
14.2.3. Example ASTM 1394 Protocol 71
14.2.4. Unsolicted Demographics Download (Viper LT) 72
14.2.4.1. Patient Fields 72
14.2.4.2. Order Fields 73
14.3. BD Phoenix 100 74
14.3.1. Configurable Options 74
14.3.1.1. General Options 74
14.3.1.2. Results Upload Options 74
14.3.1.3. Communications Setup 75
14.3.1.4. Organism and Antimicrobial Configuration Screens 76
14.3.2. Modes of Communications 76
14.3.3. Message Content 77
Field List 77
Field DescriptionsPatient Fields 78
Order Fields 78
Result Fields 79
Comment Fields 82
Terminator Field 82
14.3.4. Example ASTM 1394 Protocol 82
14.4. BACTEC FX 86
14.4.1. BACTEC FX LIS Configuration Screen 86
14.4.2. BACTEC FX Configurable Options 86
14.4.2.1. LIS ENABLE Group 86
14.4.2.2. Port Parameters Group 86
14.4.2.3. Physical Layer Group 88
14.4.2.4. LIS Options Group 88
14.4.2.5. ASTM 1381 Group 90
14.4.2.6. BD Modem Group 91
14.4.2.7. Send Button 92
14.4.2.8. Print Button 93
14.4.3. Message Content 93
14.4.3.1. Field List 93
14.4.3.2. Field Descriptions 95
14.4.4. Example ASTM 1394 Protocol 99
14.5. Innova 101
14.5.1. LIS Connection 101
14.5.2. BD ASTM Interface 102
14.5.2.1. Overview 102
14.5.2.2. Valid Characters 102
14.5.2.3. Invalid Characters 102
14.5.2.4. Message Content 103
14.5.2.5. Field List 103
14.5.2.6. Field Descriptions 105
14.5.2.7. Examples 109
14.5.2.8. CONTAINER_BC_INSTEAD_OF_ACC Barcode 111
1. Purpose
This document is intended to provide the information necessary for an LIS to exchange medical data with Becton Dickinson instruments. This document was written to include the specifications for the BD data management system as well as all future BD instruments. Any differences between instrument and data management communication are outlined in this documents appendix.
2. Scope
This document details the interface between a Laboratory Information System and Becton Dickinson instruments and the Becton Dickinson data management system. The interface uses the ASTM E_1381 protocol at the physical layer and the ASTM E_1394 protocol at the logical layer.
3. Definitions
LIS – Laboratory Information System; Computer system present in most microbiology labs responsible for collecting patient and test data.
ASTM – American Society for Testing and Materials; Committee responsible for publishing specification on communication between lab instruments and lab computer systems.
ASTM E_1381 – Protocol published by ASTM describing low-level data exchange across a serial port.
ASTM E_1394 – Protocol published by ASTM describing the logical level formatting of patient and test data.
EpiCenter – Data Management System produced by Becton Dickinson that collects microbiology test results.
LIS Interface Library – Common module responsible for implementing communication protocols. This module is used by both the BD instruments and the EpiCenter for LIS communication.
Date/Time – All times will be represented in local time unless otherwise specified.
4. Overview
Becton Dickinson (BD) has released several instruments that perform automated microbiology testing. BD has also released a data management workstation that communicates with these instruments and collects test results. It is BD’s intention to communicate these results up to the LIS, if one is present in the microbiology lab. The protocols and specifications for exchanging data between BD instruments and an LIS are described in this and subsequent sections.
It is Becton Dickinson’s expectation that an LIS vendor will implement the specifications in this document in their entirety. The proper functioning of the EpiCenter and BD instruments is dependent on the proper exchange of information with an LIS. An agreement to support LIS communication with BD instruments should be considered an agreement to implement all aspects of this specification.
4.1. Physical Cable Interface
The BD instruments and data management system communicate to the LIS using an RS232 serial port interface. The serial port interface used by the instruments uses only 3 signals in the interface cable, Transmit Data (Pin 2), Receive Data (Pin 3), and Signal Ground (Pin 5). Becton Dickinson does not believe that the inclusion of additional signals will cause any problems for the interface, however Becton Dickinson has not tested this case.
4.2. Communication Protocols
The American Society for Testing and Materials (ASTM) has published standards for how information should be exchanged between a clinical instrument and lab computer system. BD has implemented communication software to adhere to these specifications. In system configurations that include the BD data management system, the EpiCenter plays the role of the instrument.
The ASTM committee has produced standards describing both the logical level formatting and the physical exchange of data across a serial port. The logical level formatting describes how to place individual pieces of medical information into a record string, and how to combine record strings into a properly formatted message string. The physical level protocol describes how logical level message strings can be partitioned and passed across an RS232 line to an LIS.
ASTM E_1394 is the publication that outlines the logical formatting of medical information. This is the logical protocol BD uses to read/write data exchanged with the LIS. The ASTM E_1394 publication should be referred to for implementation details. A later section in this document discusses any deviations from that specification.
ASTM E_1381 is the publication that outlines the physical communication of medical information. This is the physical protocol BD uses to read/write data exchanged with the LIS. A later section in this document discusses any deviations from that specification.
4.3. System Configurations
The BD data management system is designed to communicate with several different types of BD instruments. The system configuration of these instruments is outlined in figure 1. New generation BD instruments have a TCP/IP network capability and use a local network to communicate to the EpiCenter. These results are then passed up to an LIS via the protocols described in this document.
Since the EpiCenter collects results from several different types of BD instruments, it passes several types of test results up to the LIS. This document specifies several different types of result records (per ASTM E_1394) for handling all of the various types of result data.
BD also has implemented LIS interface software in each of its new generation instruments. This allows the instruments to pass results directly up to the LIS, when an EpiCenter is not present in the lab. This configuration is illustrated in figure 2. Each of the instruments sends up the results that they acquire. The types of result messages that they produce (per ASTM E_1394) are expected to be the same as those for the EpiCenter.
Any specifications that relate only to a BD instrument are included in an appendix to this document.
4.4. EpiCenter- LIS Communication
The EpiCenter is designed to handle results from several different types of microbiology tests. BD instruments produce a varied set of results including growth and detection, ID, and AST results. In addition, manual off-line tests are often required to supplement the instrumented results. The EpiCenter is designed to manage these off-line test results as well as instrumented results.
The EpiCenter expects to upload test results, both instrumented and manual, to the LIS. The EpiCenter also expects to receive downloaded patient demographics and supplemental test information from the LIS. The test information that is exchanged with an LIS is configurable by the user, as is described in a later section of this document.
The EpiCenter is designed to use coded values for communicating data with an LIS. Most fields in the BD database are associated with an “LIS Code” that is configurable at the EpiCenter. This is a 20-character string that contains an LIS defined value that uniquely identifies that item. “LIS Codes” are used to represent patient demographic fields, tests, and test results in LIS communications. Later sections of this document outline the fields represented by LIS codes and how they are mapped in an LIS communication.
The most efficient way to implement this interface would be an exchange of configuration information between BD and the LIS vendor. The LIS vendor can provide BD with the LIS Codes necessary to translate message data. The list of tests and results currently defined in the EpiCenter database is included in Appendix B of this document. The EpiCenter can also import coded list information (Body site, Hospital Service, Ordering Physician, etc) The EpiCenter can accept this information as a text based comma separated value file or other similar format.
However, most BD instruments are not able to change their LIS codes for LIS vendor convenience. Instead they pass hard coded literal strings for test id’s, test status values, etc. The literal values passed by the instruments are restricted to fields related to that instrument. These literal strings used by the instruments are outlined in later sections of this document.
5. Workflow
The EpiCenter has a more sophisticated workflow with regard to an LIS than any instrument. Therefore this section concentrates more heavily on EpiCenter workflow. Details about BD instrument workflow are provided in an appendix in this document.
5.1. Uploaded Test Results
Each test in the EpiCenter database, both manual and instrumented, has a set of results defined for that test. Each test result has a configurable field determining if that status is reportable to the LIS. This field can be set to one of three values, always report to LIS, never report to LIS and report only when finalized. When a test result is changed to a new value, the EpiCenter checks this field to determine if the result should be uploaded to the LIS. A status configured to be “Auto Upload only when finalized” is only uploaded when the user has finalized the test. By using the “Auto Upload to LIS” field, the user can tailor the LIS interface to upload only the tests and results the LIS has defined. This configuration option also allows the EpiCenter to mimic any LIS workflow defined by a particular instrument.
Test results can be forced over these upload restrictions by first enabling the LIS Forced Upload feature of the BD EpiCenter. Then initiating a manual upload of the test results via the EpiCenter user interface, or querying for the associated specimen or test will
result in the test results being uploaded to the LIS despite the upload restrictions. The EpiCenter can be configured to enable or disable unsolicited uploads from the EpiCenter. If unsolicited uploads are enabled, then when a test is assigned an uploadable status, an upload message is generated to the LIS. If the unsolicited upload option is disabled, then test status changes do not generate upload messages. Instead, the LIS has to query the EpiCenter for test information, and only those tests with an uploadable status are included in the response message. If LIS Forced Uploads is enabled at the EpiCenter, all test results matching the query will be uploaded regardless of whether their status is uploadable.
5.2. Uploaded Isolate Results
In an attempt to characterize an organism, a lab may run more than one ID/AST test on a particular isolate. Multiple ID/AST tests may produce conflicting results for either an organism id or antibiotic susceptibility. The EpiCenter identifies when two or more tests produce conflicting results and alerts the user that they must select the appropriate results for that isolate. Once the conflict has been resolved, the EpiCenter saves the ID and AST results for that isolate separate from the test results. On the EpiCenter, these are called Isolate Level Results.
Figure 3 depicts the basic record structures within the EpiCenter.
Figure 3
The relationships and hierarchy of the finalization scheme in the EpiCenter is depicted in Figure 4.
Figure 4
The EpiCenter also includes an expert system, which evaluates ID/AST results. The input to the expert system is the results from any ID/AST test. The output of the expert system may be a different isolate AST result. Any expert system changes to AST results are maintained as a separate field, but are included as part of the isolate results. So isolate results provide both the test result and the expert system result, as described in a later section.
Isolate results may be a compilation of several ID/AST results. They are also free of any conflicts. Expert system evaluations are also available with isolate level results. For these reasons, some LIS sites may find isolate level results more valuable than test level results. The EpiCenter therefore provides a mechanism for uploading isolate level results, test level results, or both.
The configuration section of the EpiCenter allows the user to set an LIS Auto Uploadable field for isolates. This field can be set to Always “Auto Upload” to LIS, “Never Auto Upload” to LIS, “Auto Upload When Specimen Finalized” or “Auto Upload When Isolate Finalized”. This works in a similar way as test status results do for tests. When the isolate Auto Upload field is set to “Auto Upload When Specimen Finalized”, isolate level results are only uploaded once the specimen has been finalized by the user. When the isolate Auto Upload field is set to “Auto Upload When Isolate Finalized”, isolate level results are only uploaded once the isolate has been finalized by the user.
Isolate results can be forced over these upload restrictions by first enabling the LIS Forced Upload feature of the BD EpiCenter. Then initiating a manual upload of isolate results via the EpiCenter user interface, or querying for the associated specimen will result in the isolate results being uploaded to the LIS despite the upload restrictions.
Isolate results are differentiated from test results as follows:
 The Test Id in Order field (O, 5, 1, 4) is set to “ISOLATE RESULT” instead of a real test id.
 The first result record of an isolate result is always an ID record regardless of test types.
 The ID record will contain the ID test source in field (R, 4, 1, 9).
 Unlike test uploads whch have a fixed number of results, the number of results included in an isolate can flucutuate until the isolate is finalized.
 Isolate level AST result records will contain Interpreted (R, 4, 1, 4), Expert (R, 4, 1, 5) and Final (R, 4, 1, 3) AST susceptibility values whereas test level results only contain Interpreted (R, 4, 1, 4) AST susceptibility values.
 Isolate level AST result records will contain the AST test source in field (R, 4, 1, 6).
It is possible to dictate the upload order of individual antimicrobial isolate results from the BD EpiCenter through the use of Antimicrobial Ordering feature.
5.3. Unsolicited Downloads
The EpiCenter accepts unsolicited download data from the LIS containing patient demographics and supplemental test information. The fields that the EpiCenter are interested in receiving are outlined in section 8 (Message Content) of this document.
Supplemental tests and results are expected to be transmitted in the same logical placement as tests and results that are uploaded. The time/date fields for test results are used to identify LIS test results more recent than EpiCenter test results. Also a test/consumable sequence number field can be used as an identifier for a test. This allows the LIS or EpiCenter to distinguish between two tests of the same type ordered against a specimen.
An LIS Test Order can also be accepted via unsolicited download without a test/consumable sequence number. In this way the LIS can order tests on specific specimens and isolates without having to format a valid BD instrumented sequence number. These orders will be automatically associated with matching tests in the EpiCenter database as they are detected. These orders can also be manually associated with orphan tests by the EpiCenter user.
The EpiCenter does not accept edits from the LIS to key fields used to identify patients and specimen. A patient ID is the key field identifying a patient. An accession number is the key field identifying a specimen. If either of these fields is received in a download message to the EpiCenter, and these values are not matched in the EpiCenter database, they are considered new entries. Re-association of a test to a new specimen, or re-association of a specimen to a new patient, must be done at the EpiCenter.
For test order and result information, the EpiCenter evaluates the time stamps associated with that data to determine which results are more current. If the LIS information is more current, than it is used to override the EpiCenter information.
For all other data downloaded from the LIS, which does not include a time stamp, the LIS data is considered correct and overrides any EpiCenter data.
Isolate level results are not accepted in a download from the LIS.
5.4. File Import
The EpiCenter has the ability to import LIS files containing downloaded information. The user can access a screen in the EpiCenter to manually select files to import. Import files must be text files that follow the ASTM 1394 protocol, which describes the serial line interface. The import file text is sent through the same message parser that regular serial communications go through, processing the same fields that are configured for the serial interface.
Currently the EpiCenter does not export any upload information to a file.
5.5. Queries
The EpiCenter is able to generate request messages (or queries) to the LIS for information. When the EpiCenter sees a new accession number, either from user entry or an instrument, the EpiCenter uploads a request message. EpiCenter request messages, (as outlined in section 8 - Message Content), are sent for only a single specimen at a time.
Requests for information from the EpiCenter should be interpreted as requests for both patient demographics and supplemental test information for that specimen.
The EpiCenter can be configured to enable or disable queries sent to the LIS. If EpiCenter queries are enabled then the EpiCenter generates request messages to the LIS for information. If LIS queries are disabled, then no request messages are uploaded to the LIS.
The EpiCenter is also able to respond to a request generated by the LIS. The ASTM E_1394 specification describes several qualifying fields that can be included in the request for a particular set of results. The qualifying fields that the EpiCenter can process are a subset of those and are outlined in section 8 (Message Content) of this document.
The fields that are included in an uploaded response message are the same fields that are included in an n unsolicited upload message.
The EpiCenter only requests information from the LIS one specimen at a time. The EpiCenter does not use many of the qualifying fields provided for in the ASTM specifications. An example of a request from the EpiCenter is provided in section 8.2.6, outlining request record fields.
5.6. Error Reporting
The LIS interface implemented by Becton Dickinson adheres to all of the error detection and recovery outlined in the physical level protocol. The ASTM E_1381 protocol defines the use of check sums, time outs, etc., for the proper exchange of information. The rules described in that specification should be used to indicate whether or not data was properly sent between the LIS and the BD instruments. In the EpiCenter, any errors detected at the physical protocol level are also reported to the user and recorded in an error log.
Information transfers can also contain logical level formatting problems, or errors in message content. These logical errors may be determined after the physical receipt of the message has been acknowledged. The ASTM protocols do not describe a mechanism for reporting logical level errors. Instead any logical level errors are reported to the user by the EpiCenter. The Epicenter has an error log that contains any problems that occurred during the processing of an LIS download. It is the user’s responsibility to recognize and resolve logical level problems.
6. Protocol Specifications
Most of the specifications needed for implementing the communication protocols are contained in the published documents referred to Overview Section of this document. The following sections contain any clarifications or deviations to those specifications.
6.1. Physical Protocols
The BD LIS interface implements the ASTM E_1381 physical level protocol as outlined in the published specification except for the deviations described below.
6.1.1. Memory limitations
The ASTM E_1381 protocol does not place a limit on the total size of the message that can be received by the instrument. Since the EpiCenter is software resident on a standard PC, it does not have any significant restrictions on download message sizes. The EpiCenter expects to be able to handle any reasonably sized messages downloaded from the LIS. Should any download message exceed the maximum capacity of the EpiCenter, the message is rejected during transmission and an error notification is displayed to the user.
6.1.2. Unpacked Frames
Becton Dickinson interprets the ASTM E_1381 specification to imply that an intermediate frame of a message should be packed to be 240 characters in length. The only frame that should be less than 240 bytes in size is the last frame of a message.
However BD has had feedback from several LIS vendors who interpreted the ASTM E_1381 specification differently. These vendors have requested that the BD LIS interface be able to send and receive frames that contain only a single logical record. (i.e. Header, Patient, Order, Result and Terminator records all are sent in separate frames). This implies that most frames are less than 240 bytes in length. However it is still possible that a single logical record could be longer than the 240 bytes limit for a frame. In this case a single logical record is sent in several consecutive intermediate frames. All but the last frame are packed to 240 bytes and the last frame containing that logical record is less than 240 bytes.
The BD LIS interface is configurable to send either packed frames or unpacked frames, which contain a single record per frame. (The BD simulator described in a later section also has this configurable option).
6.1.3. BDMODEM
The BD LIS Interface can be configured to use the BDMODEM physical layer protocol. This customized protocol has been implemented by BD instrumemtation already out in the field. In order to maintain backwards compatibilty, the data management system supports this protocol. However, this protocol is not intended to be used for any new LIS connections. BD strongly suggests implementing the industry standard ASTM 1381 protocol for all new LIS connections. A full description of the BDMODEM protocol is provided in other BD documentation. If needed, please contact Becton Dickinson for more information.

## Table 2
6.1.3.1. Emulate BACTEC 9000 LIS Interface
The Emulate BACTEC 9000 LIS Interface option is available when BDMODEM is the selected physical protocol.
This option provides backwards compatibility when an EpiCenter is being installed where a BACTEC instrument is currently connected directly to an LIS per the BACTEC 9000 LIS Vendor Interface Specification.
6.1.3.1.1.Field Map Changes
The following field mappings are changed when the Emulate BACTEC 9000 LIS Interface option is enabled.
EpiCenter Default Value | Corresponding BDModem/Emulate 9000 Value
Test Code to O, 5, 1, 4 | Test Sequence Number to O, 5, 1, 4 Used for Vial Sequence Number 1
Isolate Source Test 1 to O, 14, 1, 1 | Isolate Source Test 1 to O, 5, 2, 4 Used for Vial Sequence Number 2
Isolate Source Test 2 to O, 14, 2, 1 | Isolate Source Test 2 to O, 5, 3, 4 Used for Vial Sequence Number 3
Specimen User Text 1 to O, 19, 4, 1 | Specimen User Text 1 to O, 5, 4, 4 Used for Vial Sequence Number 4
Specimen User Text 2 to O, 19, 5, 1 | Specimen User Text 2 to O, 5, 5, 4 Used for Vial Sequence Number 5
Result Type to R, 3, 1, 4 and enabled | Result Type to R, 3, 1, 3 and disabled
Sequence Number to R, 3, 1, 5 | Sequence Number to R, 3, 1, 4
Media Type to R, 14, 1, 2 | Media Type to R, 3, 1, 5
Instrument Type to R, 14, 1, 1 | Instrument Type to R, 14, 1, 5
Instrument Location to R, 14, 1, 5 | Instrument Location to R, 14, 1, 1
Request Test Status Code to Q, 5, 1, 2 | Request Test Status Code to Q, 5, 1, 4
Request Instrument Number Q, 5, 1, 4 | Request Instrument Number to Q, 5, 1, 2
Comment Text C, 4, 1, 1 | Comment Text C, 3, 1, 1
6.1.3.1.2.Test Status LIS Codes
Duplicate LIS Codes are permitted for BACTEC and FX tests when the Emulate BACTEC 9000 LIS Interface option is enabled.
6.2. Logical Protocols
The BD LIS interface implements the ASTM E_1394 logical level protocol as outlined in the published specification except for the deviations described below.
6.2.1. Unprocessed Records
The ASTM E_1394 logical level specification describes the use of Scientific and Manufacturer records. The BD LIS interface accepts these records in a download
message, however it does not evaluate the content of these record types. These records are extracted from any messages downloaded from the LIS. Messages generated by a BD LIS interface do not contain any of these records.
6.2.2. Download Only Fields
The ASTM E_1394 specification describes several fields that are divided into components, such as patient address field and patient name. For example, the address field is defined to have the street address, city, state, country, and zip code all separated by component delimiters. However these fields are saved as a single string in the EpiCenter database. The BD LIS interface reads those individual components from a download message and concatenate them for storage.
The BD LIS interface does parse these fields into their components for upload messages. If a concatenated field is selected for upload, the entire string is placed in the fist component of the field. For example, a Patient’s address, including city, state and zip code, are all included in the string in the Street Address field of the Patient record.
Currently the only two fields that are handled this way are Patient Name and Patient Address.
6.2.3. Communication Errors
The ASTM E_1394 protocol defines a hierarchical relationship to logical records. This protocol states that if a complete message is not successfully transmitted, then only the logical records that complete the hierarchical relationship must be resent. For example, if a message containing several patient records is not received by the instrument properly, then only the logical records from the latest Patient record on must be resent in a new message.
However the ASTM E_1381 protocol, in section 6.5.2.3, states that if a message is not received properly, the complete message is resent. The BD LIS interface follows the specification as outlined in the ASTM E_1381 physical level protocol. If a message is not transmitted successfully the BD LIS interface expects the whole message to be resent.
6.2.4. LIS Codes
The EpiCenter uses “LIS Codes” to represent many fields in an LIS communication. “LIS Codes” are values defined by the LIS that uniquely identify a field. BD uses these codes for any information that can be configured by the user. These fields include many patient demographic fields, such as Hospital Service. “LIS Codes” are also used as test ID values, and as test status values, as described in the overview section.
“LIS Codes” are used for field values instead of full names, as described in the ASTM E_1394 specification. If an LIS does not have abbreviated codes for these fields, then the “LIS Codes” can be configured to be the same as the field name text. A complete list of fields associated with an LIS code can be found in section 8 (Message Content) of this document.
Again, instruments generally do not communicate using LIS Codes. They send pre- defined literal strings in their ASTM messages. The strings used by the instruments when reporting their results are outlined in the instrument specific Appendices of this document.
LIS codes are NOT case sensitive and have a maximum length of 20 characters, except resistance markers that have a maximum length of 16 characters.
6.2.5. Blanking fields
As per the ASTM 1394 specification, the EpiCenter will accept the “” string as an indication that the value of a field should be blanked. Key fields, like Patient Id and Accession Number do not support this feature. When any coded list field in the EpiCenter, for example Hospital Service, gets blanked by the LIS, its value will change to the default value of “Unspecifed”.
The EpiCenter can exchange fields with an LIS that include a list of possible values, for example Resistance Markers. When the EpiCenter receives data in a download it will replace its entire list with the list provided in the download. And in order to blank out all entries of that list, the “” string should be placed in the first field designated for that data.
7. Configurable Options
There are several aspects of the BD LIS interface that are configurable by the user. This allows the BD instruments to tailor their connection to the local LIS. The configurable parts of the BD LIS interface are described below.
In some cases, configuration options are specific to the EpiCenter. An appendix at the end of this document lists configuration options specific to each instrument.
7.1. Port parameters
The BD LIS interface is based on serial port communication. The parameters for initializing the serial port are configurable by the user to allow for proper communication. The configurable parameters include the following fields.
Comm Port Number – acceptable values range from 1 to 4.
Baud Rate – acceptable values include 300, 600, 1200, 2400, 4800, 9600,
14400, 19200, 28800, 38400, 57600, and 115200.
Data Bits – acceptable values range from 5 to 8. Stop Bits – acceptable values include 1 and 2.
Parity – acceptable values include ODD, EVEN, and NONE.
7.2. Auto-Populating Lists
The EpiCenter has a configurable option to allow new entries detected in a downloaded message to be automatically added to the BD database. If an unrecognized “LIS Code” is found in a download field, a new entry is made to the database containing that translation code. A notification is then presented to the user suggesting that they complete the proper name, and other fields, for that list entry. This is designed to provide a mechanism for initializing the EpiCenter database with all of the relevant values needed for proper LIS communication.
Antibiotics and organisms are considered coded list fields by the EpiCenter and have an LIS translation code associated with them. For antibiotics and organisms the LIS codes are defaulted to Becton Dickinson defined abbreviations. The user has the ability to modify these codes if they do not agree with codes used by the LIS, however the Organism and Antimicrobial coded lists do not support the auto-populate feature. A listing of antibiotics and corresponding BD abbreviations can be found in Appendix C. A listing of organisms and corresponding BD abbreviations can be found in Appendix D.
7.3. Logical Protocol Parameters
The ASTM E_1394 protocol specifies four types of characters to be used to delimit fields in a record, (field, repeat, component, and escape delimiters). The user is able to configure which characters are used as delimiter characters. These characters are defaulted to the ones used in the ASTM E_1394 publication, ‘¦’, ‘\’, ‘^’, and ‘&’ respectively.
The user is able to configure whether packed or unpacked frames are uploaded to the LIS. Unpacked frames contain a single logical record per frame.
At the EpiCenter, the user is able to configure whether or not a logical record is terminated by a CR character only, or both a CR – LF character sequence.
7.4. ASTM Field Mapping
There are a number of fields contained in the EpiCenter database that need to be exchanged with the LIS. BD has made an effort to properly map these fields, per the ASTM E_1394 specification, to a field position in an appropriate record. However, some LIS vendors may interpret the meaning of these ASTM fields differently. The BD LIS interface allows the user to update the location of any field. This should allow the user to compensate for any inconsistencies between BD and LIS interfaces.
The user is able to identify each field in the BD database as an upload or download field. Fields enabled for upload messages are included in result messages passed up to the LIS. Fields enabled for download messages are parsed from any download messages from the LIS. This allows the BD LIS interface to be tailored to the specific needs of the LIS environment.
Several fields appear to be mapped into multiple locations, including test/consumable sequence number and Organism Code. These fields are mapped into the Order record because they may be needed to define a specific ordered test. These fields are also mapped into the Result record because they may be considered a result of an ID test. These fields are looked for and extracted from either location.
7.5. User Defined Fields
The EpiCenter contains 5 patient and 5 specimen fields that can be defined by the user. The EpiCenter makes no assumptions about what type of information is maintained in these fields. The user is able to use these fields to contain any data that the EpiCenter did account for with its other fields.
Three patient fields and three specimen fields are managed as coded list fields. These are fields whose value is one from a pre-configured list. For example the Specimen Type field can be set to a value from a determined list. The user is able to configure the entries listed in each field and define a label for that field.
Two more patient and 2 more specimen fields are free text fields that need not come from a pre-defined list of values. For example, these fields could be used for data such as a patient’s Social Security Number. The user can also define a label for these fields.
Since the EpiCenter has no knowledge of the type of information being maintained in these fields, it cannot assign a meaningful default mapping in an ASTM record. These fields are mapped by default to an unused ASTM location. They are also disabled by

## Table 3
default for both upload and download messages. The user can populate these fields and re-map them to an ASTM location agreed upon with the local LIS.
8. Message Content
The following sections describe which fields are exchanged with each instrument and the EpiCenter.
8.1. Field List
This section lists all of the fields by name EpiCenter can exchange with the LIS.
Each field is displayed with its default mapping in the ASTM E_1394 records. This position consists of a record type (Patient, Order, Result, Query, or Comment), a field delimiter counter, repeat delimiter counter, and component delimiter counter. The fields are grouped according to the ASTM E_1394 record type they are mapped to.
The field list also indicates if that field is defaulted as an upload (U) or download (D) field, or both (U/D). A blank column indicates that the field is not to be exchanged with an LIS. Fields not configured to be exchanged with the LIS are ignored in a download message.
The first two fields of every record include the Record Type indicator and a record index value. These fields are considered part of every record but are not listed in these tables. The following section describes these fields in more detail, including the expected values for each field.
Result records have data fields with multiple definitions, such as Result Data 1, Result Data 2, etc. The data contained in these fields depends on the type of results being reported. ID tests may have organism data in those fields, while AST tests have antibiotic data in those fields. The following sections describe in more detail what data can be expected in a result record.
Header Record Field Name | ASTM Pos. | Direction
Sender Name | H, 5, 1, 1 | U
Version Number | H, 13, 1, 1 | U
Message Date/Time | H 14, 1, 1 | U
Patient Record Field Name | ASTM Pos. | Direction
Patient ID | P, 4, 1, 1 | U/D
Patient Name (Last) | P, 6, 1, 1 | U/D
Patient Name (First) | P, 6, 1, 2 | D
Patient Name (Middle) | P, 6, 1, 3 | D
Patient Name (Suffix) | P, 6, 1, 4 | D
Patient Name (Title) | P, 6, 1, 5 | D
Date of Birth | P, 8, 1, 1 | U/D
Patient Sex | P, 9, 1, 1 | U/D
Address (Street) | P, 11, 1, 1 | U/D
Address (City) | P, 11, 1, 2 | D
Address (State) | P, 11, 1, 3 | D
Address (Zip Code) | P, 11, 1, 4 | D
Address (Country) | P, 11, 1, 5 | D
Patient Phone Number | P, 13, 1, 1 | U/D
Admitting Physician | P, 14, 1, 1 | U/D
Patient User Field 1 | P, 15, 1, 1
Patient User Field 2 | P, 15, 2, 1
Patient User Field 3 | P, 15, 3, 1
Patient User Field 4 | P, 15, 4, 1
Patient User Field 5 | P, 15, 5, 1
Patient Diagnosis | P, 19, 1, 1 | U/D
Patient Therapy (1-5) | P, 20, 1-5,1 | U/D
Admit Date/Time | P, 24, 1, 1 | U/D
Room Number | P, 26, 1, 1 | U/D
Hospital Service | P, 33, 1, 1 | U/D
Hospital Client | P, 34, 1, 1 | U/D

## Table 4
Order Record Field Name | ASTM Pos. | Direction
Accession Number | O, 3, 1, 1 | U/D
Isolate Number | O, 3, 1, 2 | U/D
Organism | O, 3, 1, 3 | U/D
Exclude Isolate from Statistics | O, 3, 1, 4 | U/D
Test ID | O, 5, 1, 4 | U/D
Test/Consumable Sequence Number | O, 5, 1, 5 | U/D
Priority | O, 6, 1, 1 | U/D
Collection Date/Time | O, 8, 1, 1 | U/D
Collected By | O, 11, 1, 1 | U/D
Received By | O, 11, 1, 2 | U/D
Specimen Action Code | O, 12, 1, 1 | U
Isolate Source Test (1-3) | O, 14,1-3,1 | U/D
Isolate Source Test Start Time (1-3) | O, 14,1-3,2 | U/D
Receipt Date/Time | O, 15, 1, 1 | U/D
Specimen Type | O, 16, 1, 1 | U/D
Body Site | O, 16, 1, 2 | U/D
Ordering Physician | O, 17, 1, 1 | U/D
Ordering Physician Phone | O, 18, 1, 1 | U
Ordering Physician Fax | O, 18, 1, 2 | U
Ordering Physician Pager | O, 18, 1, 3 | U
Specimen User Field 1 | O, 19, 1, 1
Specimen User Field 2 | O, 19, 2, 1
Specimen User Field 3 | O, 19, 3, 1
Specimen User Field 4 | O, 19, 4, 1
Specimen User Field 5 | O, 19, 5, 1
Finalized Date/Time* | O, 23, 1, 1 | U/D
Specimen Reimbursement Value | O, 24, 1, 1 | U
Test Reimbursement Value | O, 24, 2, 1 | U
Isolate Classification/Nosocomial | O, 29, 1, 1 | U/D
Result Record Field Name | ASTM Pos. | Direction
Result Type Code | R, 3, 1, 4 | U/D
Test/Consumable Sequence Number | R, 3, 1, 5 | U/D
Antibiotic | R, 3, 1, 6 | U/D
Antibiotic Concentration | R, 3, 1, 7 | U/D
Antibiotic Concentration Units | R, 3, 1, 8 | U/D
Test Status | R, 4, 1, 1 | U/D
Result Data 1
BACTEC MGIT 960 Growth Units | R, 4, 1, 2 | U

## Table 5
Result Record Field Name |  | ASTM Pos. |  | Direction
ProbeTecET/Viper XTR Alg.Results |  | R, 4, 1, 2
AST MIC for AST MIC test |  | R, 4, 1, 2 |  | U/D
AST Diameter for AST Dia. test |  | R, 4, 1, 2 |  | U/D
Organism ID for ID test |  | R, 4, 1, 2 |  | U/D
Result Data 2
AST susceptibility (Final) |  | R, 4, 1, 3 |  | U/D
Profile Number for ID Test |  | R, 4, 1, 3 |  | U/D
Profile Number for Isolate ID |  | R, 4, 1, 3 |  | U/D
Result Data 3
AST susceptibility (Interpreted) |  | R, 4, 1, 4 |  | U
Resistance Marker 1 |  | R, 4, 1, 4 |  | U
Result Data 4
AST susceptibility (Expert) |  | R, 4, 1, 5 |  | U
Resistance Marker 2 |  | R, 4, 1, 5 |  | U
Result Data 5
AST Source test |  | R, 4, 1, 6 |  | U
Resistance Marker 3 |  | R, 4, 1, 6 |  | U
Result Data 6
Resistance Marker 4 |  | R, 4, 1, 7 |  | U
Result Data 7
Resistance Marker 5 |  | R, 4, 1, 8 |  | U
Result Data 8
ID Source Test |  | R, 4, 1, 9 |  | U
Preliminary/Final Status |  | R, 9, 1, 1 |  | U/D
Test Start Date/Time |  | R, 12, 1, 1 |  | U/D
Result/Status Date/Time |  | R, 13, 1, 1 |  | U/D
Test Complete Date/Time |  | R, 13, 2, 1 |  | U
Media Setup Date/Time |  | R, 13, 1, 2 |  | U
Instrument Type |  | R, 14, 1, 1 |  | U
Media/Assay Type |  | R, 14, 1, 2 |  | U
Protocol Length |  | R, 14, 1, 3 |  | U
Instrument Number |  | R, 14, 1, 4 |  | U
Instrument Location |  | R, 14, 1, 5 |  | U
Protocol Name |  | R, 14, 1, 8 |  | U
Additional Result Quantity (1-5) |  | R, 15, 1-5,1 |  | U/D
Additional Result (1-5) |  | R, 15,1-5,2 |  | U/D
Comment Record Field Name |  | ASTM Pos. |  | Direction
Comment Text |  | C, 4, 1, 1 |  | U/D
Comment Type |  | C, 5, 1, 1 |  | U/D
Request Record Field Name | ASTM Pos. |  | Direction
Request Starting Patient Id | Q, 3, 1, 1 |  | D
Request Starting Accession Number | Q, 3, 1, 2 |  | U/D
Request Starting Sequence Number | Q, 3, 1, 3 |  | D
Request Ending Patient Id | Q, 4, 1, 1 |  | D
Request Ending Accession Number | Q, 4, 1, 2 |  | D
Request Ending Sequence Number | Q, 4, 1, 3 |  | D
Request Test Id | Q, 5, 1, 1 |  | D
Request Test Status | Q, 5, 1, 2 |  | D
Request Instrument Type | Q, 5, 1, 3 |  | D
Request Instrument Number | Q, 5, 1, 4 |  | D
Request Result Qualifier | Q, 5, 1, 5 |  | D
Request Time Qualifier | Q, 6, 1, 1 |  | D
Starting Date/Time | Q, 7, 1, 1 |  | D
Ending Date/Time | Q, 8, 1, 1 |  | D
Request Information Status Code | Q, 13, 1, 1 |  | U/D

## Table 6
Terminator Record Field Name | ASTM Pos. | Direction
Termination Code | L, 3, 1, 1 | U/D
8.2. Field Descriptions
This section provides detailed information for each field listed in the previous section.
8.2.1. Header Fields
The header records that start each logical message are the same from all BD instruments. An example record is provided below. For examples of Header records used in the context of a message, please refer to Appendix A of this document.
H¦\^&¦ ¦ ¦Becton Dickinson¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦V1.00¦19981019184200
Delimiter Fields (H, 2, 1, 1) – These single characters can be used to process the remainder of the message. These characters denote the field, repeat, component and escape delimiters as described in the ASTM E_1394 specification. These fields are defaulted to “¦”, “\”, “^”, and “&” respectively, as shown above.
Sender Name (H, 5, 1, 1) – Messages coming from a Becton Dickinson instrument have the BD identifier in this field as shown above.
Version Number (H, 13, 1, 1) – This version number represents the version of the BD LIS interface used for communications.
Message Date/Time (H, 14, 1, 1) – BD includes the current time and date when constructing his message, formatted as described in the ASTM E_1394 specification in section 6.6.2.
8.2.2. Patient Fields
A full patient record is shown below. The fields used in this example are referred to in the field explanations. For examples of Patient records used in the context of a message, please refer to Appendix A of this document.
P¦1¦ ¦PatId123¦ ¦Doe^John^R^Jr.^Dr.¦ ¦19651029¦M¦ ¦2 Main St.^Baltimore^ MD^21211^USA¦ ¦(410) 316 - 4000¦JSMITH¦ ¦ ¦
¦ ¦PNEU¦P\AM\AMX¦ ¦ ¦ ¦19981015120000¦ ¦324¦ ¦ ¦ ¦ ¦ ¦ ¦ER¦St. Josephs Hospital
Patient ID (P, 4, 1, 1) - Identifier that uniquely identifies a patient. This is a required field for patient demographic downloads. This field can be up to 16 characters long. (e.g. PatId123)
Patient Name (P, 6, 1, 1-5) – Patient name is divided into 5 components in the ASTM E_1394 specification. The BD EpiCenter stores these fields as a single string in its database. Therefore the BD LIS Interface concatenates these fields from a download message. The patient name can be up to 40 characters long. If these fields are selected for upload on the EpiCenter the whole name is placed in the first component of this field (P, 6, 1, 1). Patient name shown in the sample record is Dr. John R Doe Jr.
Date of Birth (P, 8, 1, 1) – Patient date of birth formatted as described in ASTM E_1394 section 6.6.2. The date shown in the sample record is Oct. 29, 1965.
Patient Sex (P, 9, 1, 1) – There are 3 predefined values accepted for this field in the BD EpiCenter database, Male, Female, and Unspecified. If the value does not match the code for either Male or Female then the sex is set to Unspecified. (This is an LIS coded field for the EpiCenter).
Patient Address (P, 11, 1, 1-5) – The ASTM E_1394 specification suggests that the address field be divided into 5 separate components. The EpiCenter concatenates all 5 address fields into one before saving it in the database. The patient address can be up to 255 characters long. If this field is selected for upload on the EpiCenter, the whole address is placed in the first component of this field (P, 11, 1, 1).
Patient Phone Number (P, 13, 1, 1) – This field is defined as a free text field that can be up to 25 characters long.
Admitting Physician (P, 14, 1, 1) – The physician admitting the patient into the hospital. In the example above “JSMITH” could be used as an LIS defined abbreviation for Dr.
Jane T. Smith, who may work in the hospital. (This is an LIS coded field for the EpiCenter).
Patient Diagnosis (P, 19, 1, 1) – The diagnosis assigned to the patient for the current visit. In the example above “PNEU” could be used as an LIS defined abbreviation for Pneumonia. (This is an LIS coded field for the EpiCenter).
Patient Therapy (P, 20, 1-5, 1) – This field contains up to 5 antibiotics that the patient may be receiving at the time the specimen is being processed. A repeat delimiter as per ASTM E_1394 specifications separates each antibiotic. The above example indicates that the patient is currently receiving Penicillin (P), Ampicillin (AM), and Amoxicillin (AMX). (This is an LIS coded field for the EpiCenter).
Admission Date/Time (P, 24, 1, 1) – This is the time and date that the patient was admitted to the hospital. This field is formatted as specified in the ASTM E_1394 specification, in section 6.6.2. The time portion of this field is assumed to be local time. The optional time zone is deemed unnecessary and is not evaluated.
Note: If the Date/Time string is less than fourteen-characters, the time shall be set to 12:00:00 AM.
Room Number (P, 26, 1, 1) – This is a text field containing the patient’s room number and can be up to 10 characters long.
Hospital Service (P, 33, 1, 1) – The ward or hospital service involved in caring for this patient. (This is an LIS coded field for the EpiCenter). (e.g. ER may represent the Emergency Room).
Hospital Client (P, 34, 1, 1) – This field is used by reference laboratories that accept specimens from client hospitals. (This is an LIS coded field for the EpiCenter).
User Defined Patient Fields 1-3 (P, 15, 1-3, 1) – In the EpiCenter, the user has the ability to configure the values for 3 coded lists related to patient fields. These fields are mapped to unused fields in the patient record by default. The user may fill these fields and remap them to a field agreed upon by the local LIS. (These are LIS coded fields for the EpiCenter).
User Defined Patient Fields 4-5 (P, 15, 4-5, 1) – In the EpiCenter, the user has 2 free text fields that can hold any alpha-numeric information. These fields are mapped to unused fields in the patient record by default. The user may fill these fields and remap them to a field agreed upon by the local LIS.
8.2.3. Order Fields
A full order record is shown below to provide examples for the content of each field. For examples of Order records used in the context of a message, please refer to Appendix A of this document.
O¦1¦Acc123^1^MYCBTUB ¦ ¦ ^ ^ ^ MGIT_960_GND ^Seq123¦
¦ ¦19981019023300¦ ¦ ¦SJB^MMF¦A¦ ¦ ¦19981019045200
¦Blood^Arm¦MJones¦(410) 555 – 1234^(410) 555 – 9876^(410) 555 – 7777¦ ¦ ¦ ¦ ¦19981020053400¦62¦ ¦O¦ ¦ ¦Nos
Accession Number (O, 3, 1, 1) – The unique alphanumeric string that identifies a specimen. This field can be up to 20 characters long. (e.g. Acc123) This is a required field for processing specimen and test information.
Isolate Number (O, 3, 1, 2) – The number associated with an isolate that the test should be performed on.. This field should only be used when ordering isolate based tests, otherwise this field should be left blank. Isolate numbers can range from 1 to 20.
Organism (O, 3, 1, 3) – The organism assigned to a particular isolate. The organism code shown in the example above, MYCBTUB, is the code for Mycobacterium tuberculosis. (This is an LIS coded field for the EpiCenter).
If an organism is indicated when ordering a BD Phoenix 100 ID/AST combo panel, then the ID portion of the panel will automatically be disabled. As a result, only the AST portion of the panel will report results.
Exclude Isolate from Statistics (O, 3, 1, 4) – This field will explicitly indicate if this isolate should be included or excluded from statistic reporting. If this field contains an “I”, then the isolate will be included in statistical reporting. If this field contains an “E”, then the isolate will be excluded from statistical reporting. If this field is left empty, then the EpiCenter will use its internal logic to make this decision.
Test Id (O, 5, 1, 4) – The unique code used to represent a particular test. This is an LIS coded field for the EpiCenter. For BD instrumented tests, the test ID is the pre-defined test name provided in appendix B. (e.g. MGIT_960_GND)
For isolate level results coming from the EpiCenter, this field will contain the string “ISOLATE RESULT”.
As noted in Appendix B, tests common to the BACTEC 9000 and BACTEC FX have the same LIS code.
Test/Consumable Sequence Number (O, 5, 1, 5) – The value that identifies a test ordered. For tests originating from an instrument, this field is set with the consumable id value.
For manual tests ordered at the EpiCenter, this field contains any test id entered by the user at test order time. In downloads from the LIS, this field is used to distinguish between multiple tests (of the same type) ordered against a specimen. If this field is blank, the EpiCenter assumes that this is an LIS Test Order and creates a blank test in the EpiCenter database. These orders will be automatically associated with matching tests in the EpiCenter database as they are detected. These orders can also be manually associated with orphan tests by the EpiCenter user.
This value can be up to 17 characters long.
Priority (O, 6, 1, 1) – This field is defined in the ASTM E_1394 specification. Currently only the values ‘A’ for Critical, ‘R’ for Normal and blank for default normal or no change are supported.
Collection Date/Time (O, 8, 1, 1) – Date and time that the specimen was collected from the patient, formatted as described in ASTM E_1394 section 6.6.2. The date/time shown in the sample record is 2:33 am on Oct. 19, 1998.
Note: If the Date/Time string is less than fourteen-characters, the time shall be set to 12:00:00 AM.
Collected By (O, 11, 1, 1) – The hospital personnel that collected the specimen from the patient. The example above shows that the specimen was collected by SJB which could be the initials for Sarah J Baker. (This is an LIS coded field for the EpiCenter).
Received By (O, 11, 1, 2) – The lab technician that received the specimen. The example above shows that the specimen was received by MMF, which could be the initials for Mike M Ferris. (This is an LIS coded field for the EpiCenter).
Specimen Action Code (O, 12, 1, 1) – This field is defined in the ASTM E_1394 specification. Currently only the value Q is supported, for QC tests. Normal tests leave this field blank.
Isolate Source Test (1-3) (O, 14, 1-3, 1) – The isolate source test is the specimen level test that produced the isolate. Often growth and detection tests are sub-cultured to produce several isolates. This field will contain test ids for up to 3 tests that yielded an isolate. This field only applies to Order records containing isolate information.
Isolate Source Test Start Time (1-3) (O, 14, 1-3, 2) – This is the start date and time of the isolate source test described above. It is possible that two of the same test type will be run on a specimen. This field is used to uniquely identify the isolate source test. This field only applies to Order records containing isolate information. This field is formatted as described in ASTM E_1394 section 6.6.2.
Receipt Date/Time (O, 15, 1, 1) – Date and time that the specimen was received into the lab, formatted as described in ASTM E_1394 section 6.6.2. The date/time shown in the sample record is 2:33 am on Oct. 19, 1998.
Note: If the Date/Time string is less than fourteen-characters, the time shall be set to 12:00:00 AM.
Specimen Type (O, 16, 1, 1) – The type of specimen collected from the patient. (This is an LIS coded field for the EpiCenter). (e.g. Blood)
Body Site (O, 16, 1, 2) – The body area that the specimen was taken from. (This is an LIS coded field for the EpiCenter). (e.g. Arm)
Ordering Physician (O, 17, 1, 1) – The physician that ordered this test. (This is an LIS coded field for the EpiCenter). (e.g. MJones could represent Mark Jones).
Ordering Physician Phone Numbers (O, 18, 1, 1-3) – The phone numbers that can be used to reach the ordering physician. Three numbers are accepted in this field, phone number, fax number and pager number. A component delimiter separates each phone number. These fields are free text fields that can be up to 25 characters in length.
User Defined Specimen Fields 1-3 (O, 19, 1-3, 1) – In the EpiCenter, the user has the ability to configure the values for 3 coded lists related to specimen fields. These fields are mapped to unused fields in the patient record by default. The user may fill these fields and remap them to a field agreed upon by the local LIS. (This is an LIS coded field for the EpiCenter).
User Defined Specimen Fields 4-5 (O, 19, 4-5, 1) – In the EpiCenter, the user has 2 free text fields that can hold any alphanumeric information. These fields are mapped to unused fields in the order record by default. The user may fill these fields and remap them to a field agreed upon by the local LIS.
Finalized Date/Time (O, 23, 1, 1) – The ASTM specification describes this field as containing the time the results were last modified. The EpiCenter places this information in the result record to be evaluated for each result. This field contains the date/time, formatted as described in ASTM E_1394 section 6.6.2, that the results for this test were finalized.. This field is blank if the test results are not yet finalized. The time shown in the example record is 5:34 am on Oct 20, 1998.
For isolate level results, this field contains the date/time the specimen was finalized.
Specimen Reimbursement Value (O, 24, 1, 1) – The value used for insurance reimbursement based on the specimen type. This number can be between 0 and 10,000. The value of 62 is used in the example record above.
Test Reimbursement Value (O, 24, 2, 1) – The value used for insurance reimbursement based on this test type. This number can be between 0 and 10,000. The example record above does not have a test reimbursement value.
Report Type (O, 26, 1, 1) – This field is defined in the ASTM E_1394 specification. When a cancelled order is sent back to the LIS the report type is ‘X’.
Isolate Classification/Nosocomial (O, 29, 1, 1) – This field contains the EpiCenter LIS code for the isolate classification. The EpiCenter has a pre-defined list of classifications for an isolate. The complete list is as follows (Significant/Nosocomial, Significant/Community Acquired, Significant/Unknown, Contaminate, Not Determined). If a downloaded code is not recognized, then the Isolate Classification is set as Significant/Unknown.
8.2.4. Result Fields
The EpiCenter uploads results from several different types of tests. This causes several different types of result records to be passed to the LIS. The information contained in the generic “Data Fields” depends on the type of result being reported. AST Result records contain antibiotic and susceptibility information. ID Result records contain organism and Resistance Marker information. Growth and Detection result records contain a Positive or Negative status. Below are several different examples of result records, each one passing up a different type of result.
For AST sets, several result records may be sent to the LIS in a single message. Each result record would report the results for a single drug in the AST set. For ID/AST combo tests, an ID result record is sent as the first result record, and the AST result records follow. An example AST set result message is shown in the Appendix A of this document.
Isolate results are also exchanged using these records. Isolate results are exchanged using the same format as an ID/AST combo tests with an ID record first, then followed by several AST result records. Since isolate results are not associated with a particular test, the test code field in the Order record is replaced with the literal string “ISOLATE_RESULT”.
For examples of Results records used in the context of a message, please refer to Appendix A of this document.
BACTEC MGIT 960 growth and detection test level result example
R¦1¦ ^ ^ ^GND_MGIT^430100001234¦INST_POSITIVE
^87¦ ¦ ¦ ¦ ¦P¦ ¦ ¦19981019153400¦19981020145000¦ MGIT960^^42^3^B/A12
BACTEC MGIT 960 AST test level result example:
R¦1¦ ^ ^ ^AST_MGIT^439400005678^P^0.5^ug/ml¦ INST_COMPLETE^105^^S¦ ¦ ¦ ¦ ¦P¦ ¦ ¦19981019153400¦ 19981020145000¦MGIT960^^42^3^ B/A12
Phoenix AST MIC test level result example:
R¦1¦ ^ ^ ^AST_MIC^429530000002^P¦ INST_COMPLETE^0.5^^S¦ ¦ ¦ ¦ ¦ F¦ ¦ ¦19981019153400¦ 19981020145000
Diameter based AST test level result example:
R¦1¦^^^AST_DIA^Seq123^P^0.5¦COMPLETE^15^^S¦ ¦ ¦
¦ ¦F¦ ¦ ¦19981019153400¦ 19981020145000
Identification test level result example:
R¦1¦ ^ ^
^ID^Seq123¦Complete^MYCBTUB^45678^RM_VRE¦ ¦ ¦
¦ ¦F¦ ¦ ¦19981019153400¦19981020145000
For all other tests:
R¦1¦ ^ ^ ^OTHER^Seq123¦Complete¦ ¦ ¦ ¦ ¦F¦ ¦
¦19981019153400¦19981020145000¦
¦Many^GPOS_Bac\Few^GNEG_Rods
Isolate level result example:
R¦1¦ ^^^AST^^P^100.0^ug/mL¦
^^R^R^^MGIT_960_AST92¦ ¦ ¦ ¦ ¦F
“Universal Test ID” ASTM E_1394 Result field 3
Result ID Code (R, 3, 1, 4) – This is a BD defined code that indicates the type of information being exchanged in the result record. The interpretation of several fields may vary depending on which type of result is being passed. The possible values for this field include the following.
GND – indicates results for an FX, BACTEC 9000 or generic growth and detection test.
GND_MGIT – indicates results for a BACTEC MGIT 960growth and detection test.
GND_PROBETEC – indicates results for a BDProbeTec ET test.
AST – indicates results for a generic AST test not included in the other AST categories.
AST_MGIT – indicates results for a BACTEC MGIT 960 Antibiotic Susceptibility Test.
AST_MIC – indicatesresults for an MIC based Antibiotic Susceptibility Test.
AST_DIA – indicates results for a diameter based Antibiotic Susceptibility Test. (e.g. Kirby Bauer tests).
ID – indicates identification results.
STREAK – indicates Innova Protocol results
OTHER – indicates test results not included in the above categories.
Test/Consumable Sequence Number (R, 3, 1, 5) – The value that identifies the media/consumable used for generating these results. This field may be blank for manual test results where a sequence number is not known. This value can be up to 17 characters long. The example records above show the results being reported for sequence number “Seq123”.
A test/consumable sequence number is required for all instrumented test results downloaded to the EpiCenter. For manual or supplemental test results downloaded to the EpiCenter this field is optional. If no sequence number is included with the results the EpiCenter searches on the current access number for any test matching the test code provided in the order record. If a match is found, the EpiCenter updates the results for that test. If no test match is found, then the EpiCenter orders a new test. Be aware that unwanted consequences could occur if more than one of a particular test type are ordered for an access number and test sequence numbers are not used.
Antibiotic (R, 3, 1, 6) (AST Results) – The antibiotic tested in this AST test. The antibiotic code shown in the above AST examples is “P” for Penicillin. (This is an LIS coded field for the EpiCenter).
Concentration (R, 3, 1, 7) (AST Results) – Several AST tests, such as Kirby Bauer tests, are done with antibiotics at a specific concentration. For these tests the antibiotic concentration is needed to define the test that results are being provided for. When no concentration units are provided it is assumed to be in ug/ml. This field can be up to 15 characters long. The antibiotic concentration shown in the AST (MGIT960) example above is “0.5”.
Concentration Units (R, 3, 1, 8) (AST Results) – The BACTEC MGIT 960 results include the concentration units of the concentration provided in the previous field. This field can be up 15 characters long. The concentration units shown in the AST example above is “ug/ml” for micrograms per milliliter. This is also assumed to be the default if no other value is provided.
“Data or Measurement Values” ASTM E_1394 Result field 4
The contents of the Result Data fields vary depending on the type of test result being reported, as described below.
Test Status (R, 4, 1, 1) – The status value for a particular test. This field is filled for all test results. This is an LIS coded field for the EpiCenter. For BD instrumented tests, the status code is one of the pre-defined values listed in Appendix B of this document. The GND example record shown above has a test status value of INST_POSITIVE.
Result Data Field 1
Growth Units (R, 4, 1, 2) (BACTEC MGIT 960 Results Only) – See Appendix E for instrument specific information. The Positive GND tube listed above indicates 87 growth units. The AST example record shows a growth value of 105. This field is never more than 5 characters long.
ProbeTecET Algorithmic Result (R, 4, 1, 2) (BDProbeTec ET/Viper XTR Result Only) – See Appendix E for instrument specific information. (future)
Minimum Inhibitory Concentration (R, 4, 1, 2) (MIC based AST Result) – For AST tests that produce MIC values, this field contains the minimum antibiotic concentration that inhibits growth of the tested organism. The example above for an MIC based AST test contains an MIC value of “0.5”. This field can be up to 20 characters long and is assumed to contain a number in ug/ml.
This field may contain signed MIC values such as “<=4” or “>8”, MIC values for compound antimicrobials such as “0.5/4”or “<=0.5/16” as well as the following:
“?” - Ongoing
“C” - Antimicrobial is Rapid Completed (For nonQC tests in EpiCenter V4 and later)
“X” - Error
Diameter (R, 4, 1, 2) (Diameter based AST Results) – For Kirby Bauer type tests, this field contains the diameter of the inhibited growth region caused by the antibiotic. The diameter based AST example above shows a diameter value of 15. This number is assumed to be in millimeters.
Organism (R, 4, 1, 2) (ID Results Only) – The organism assigned to a particular isolate. The organism code shown in the example above, MYCBTUB, is the code for Mycobacterium tuberculosis. (This is an LIS coded field for the EpiCenter).
Result Data Field 2
AST Susceptibility, Final (R, 4, 1, 3) (Isolate AST results only) – This is a code that indicates is the organism is susceptible to the antibiotic specified in this record. The acceptable values for this field are S, I, R, N, and X, representing susceptible, intermediate, resistant, not-susceptible and error respectively. This field is only uploaded by the EpiCenter as part of an isolate level result (see earlier section). The isolate AST example record shows that the user decided to override the expert system results and say that the organism was susceptible to the antibiotic, indicated by “S” in the Final field.
Profile Number (R, 4, 1, 3) (ID results and Isolate ID results) – This is a text field that holds the profile number associated with an ID test. This value can be up to 16 characters long. The ID result record example above indicates a profile number of 45678.
Result Data Field 3
AST Susceptibility, Interpreted (R, 4, 1, 4) ( All AST results) – This is a code that indicates is the organism is susceptible to the antibiotic specified in this record. The acceptable values for this field are S, I, R, N, and X, representing susceptible, intermediate, resistant, not-susceptible and error respectively. This field contains the AST interpretation from the instrument or user before being processed by the Expert System. The isolate AST example record shows that the organism has intermediate susceptibility to the specified antibiotic, indicated by “I”.
Resistance Marker 1 (R, 4, 1, 4) (Isolate ID results) – This represents a Resistance Marker associated with the organism identified in the ID result record. The EpiCenter generates this data as part of the expert system and are only included for Isolate uploads. The first Resistance Marker in the ID example above is RM_VRE, which may be an LIS code used for Vancomycin resistant Enterococci. (This is an LIS coded field for the EpiCenter).
Result Data Field 4
AST Susceptibility, Expert (R, 4, 1, 5) (Isolate AST results only) – This is a code that indicates if the organism is susceptible to the antibiotic specified in this record. The acceptable values for this field are S, I, R, N, and X, representing susceptible, intermediate, resistant, not-susceptible and error respectively. This field is only uploaded by the EpiCenter as part of an isolate level result (see earlier section). The BD Phoenix instrument and the EpiCenter have an expert system that can produce a secondary evaluation of susceptibility. The isolate AST example record shows that the expert system decided that the organism is resistant to the specified antibiotic, indicated by “R”.
Resistance Marker 2 (R, 4, 1, 5) (Isolate ID results) – This represents a Resistance Marker associated with the organism identified in the ID result record. The EpiCenter generates this data as part of the expert system and are only included for Isolate uploads. The example record above contains only one Resistance Marker. (This is an LIS coded field for the EpiCenter).
Result Data Field 5
AST Source Test (R, 4, 1, 6) (Isolate AST results) – This represents the Test Id for the source test that included this AST result. This is included because Isolate results may be compiled from several different AST tests. The example Isolate upload record above contains the Test Id of MGIT_960_AST. (This is an LIS coded field for the EpiCenter).
It is possible that there can be results on the isolate upload that have been inferred by the EpiCenter BDXpert system and are not associated with a test. When this condition occurs, EpiCenter uploads a “blank”, or an empty field in the AST Source Test field(R, 4, 1, 6) as well as the MIC field(R, 4, 1, 2).
Results example with AST Source Test: R¦2¦^^^AST^^CTX¦^<=1^S^S^^NMICID4 ¦ ¦ ¦ ¦ ¦F Inferred results example without AST Source Test:
R¦3¦^^^AST^^CZ¦^^R^^R¦ ¦ ¦ ¦ ¦F
Resistance Marker 3 (R, 4, 1, 6) (Isolate ID results) – This represents a Resistance Marker associated with the organism identified in the ID result record. The EpiCenter generates this data as part of the expert system and are only included for Isolate uploads.
The example record above contains only one Resistance Marker. (This is an LIS coded field for the EpiCenter).
Result Data Field 6
Resistance Marker 4 (R, 4, 1, 7) (Isolate ID results) – This represents a Resistance Marker associated with the organism identified in the ID result record. The EpiCenter generates this data as part of the expert system and are only included for Isolate uploads. The example record above contains only one Resistance Marker. (This is an LIS coded field for the EpiCenter).
Result Data Field 7
Resistance Marker 5 (R, 4, 1, 8) (Isolate ID results) – This represents a Resistance Marker associated with the organism identified in the ID result record. The EpiCenter generates this data as part of the expert system and are only included for Isolate uploads. The example record above contains only one Resistance Marker. (This is an LIS coded field for the EpiCenter).
Result Data Field 8
ID Source Test (R, 4, 1, 9) (Isolate ID results) – This represents the Test Id for the source test that produced this ID result. (This is an LIS coded field for the EpiCenter).Preliminary/Final Status (R, 9, 1, 1) – This field contains either a “P” or an “F” as defined in the ASTM E_1394 specification. The result status shown in the examples above are final. For isolate level results, this indicates the status of the isolate.
Start Date/Time (R, 12, 1, 1) – This is the date and time that the test was first started or entered into an instrument. This field is formatted as described in the ASTM E_1394 specification in section 6.6.2. The test start time shown above is 3:34 pm on Oct 19, 1998. This field has special significance for BACTEC MGIT 960 AST tests. Please read the instrument specific section for more details.
Result/Status Date/Time (R, 13, 1, 1) – This is the date and time that the test received the status being reported in this record. This is a required field and is used to identify when results downloaded from the LIS are more recent than those contained in the EpiCenter. This field is formatted as described in the ASTM E_1394 specification in section 6.6.2. The test status time shown above is 2:50 p.m. on Oct 20, 1998.
Test Complete Date/Time (R, 13, 2, 1) – This is the date and time that an instrumented test finished its results. This is different than the Result Date/Time in that this time will not change if the user assigns a new status to a test. For a growth and detection result record, this field will hold the Positivity Date/Time. For an ID result record, this field will hold the time that the Id portion of a BD Phoenix 100 panel completed. For an AST result from a BD Phoenix 100 panel, this field will hold the time that a valid MIC was produced. This field is formatted as described in the ASTM E_1394 specification in section 6.6.2.
Media Setup Date/Time (R, 13, 1, 2) – For Innova tests this field indicates the time that the specimen was successfully streaked. If the specimen was not successfully streaked, this field will be blank.
Instrument Type (R, 14, 1, 1) (Instrumented Results Only) – This field indicates which BD instrument produced the result. This field is filled with one of several BD defined values listed below.
“BT9000” – BACTEC 9000 series instrument “MGIT960” – BACTEC MGIT 960 instrument “PHOENIX” – BD ID/AST instrument
“PROBETEC” – BDProbeTec ET instrument
“BACTECFX” - FX instrument
“INNOVA” – Innova instrument
Media Type (R, 14, 1, 2) (Instrumented Results Only) – This field indicates the type of media or consumable used to perform this test. See Appendix E for instrument specific information.
Protocol Length (R, 14, 1, 3) (Instrumented Results Only) – this field indicates the length of time (in hours or days) required to complete the test. For growth and detection tests, this number represents the number of days before a tube can be considered negative.
Instrument Number (R, 14, 1, 4) (Instrumented Results Only) – This is the user number assigned to the instrument that ran the test. This value can be between 1 and 99.
Instrument Location (R, 14, 1, 5) (Instrumented Results Only) – This value indicates the position inside of the instrument where the test was performed.
For BACTEC MGIT 960 tests this value indicates a drawer number, row number and well number. In the GND example above, the location B/A12 indicates that the test was performed in drawer B, row A well number 12.
For Innova tests this value indicates a drawer number, canoe number, and position in canoe and is represented as 3-1-8.
BDProbeTec ET QC Type (R, 14, 1, 6) (BDProbeTec ET Result Only) – See Appendix E for instrument specific information.
Test Protocol (R, 14, 1, 8) - This value indicates the name of the protocol with which the specimen was processed on the Innova instrument.
BDProbeTec ET QC Kit Lot Number (R, 14, 1, 7) (BDProbeTec ET Result Only) – See Appendix E for instrument specific information.
Additional Results Quantity 1-5 (R, 15, 1-5, 1) (EpiCenter Additional Results Only) – This field contains a text value that associates a quantity to the additional result reported in the next ASTM component field. The quantities Many and Few are used in the example record for other tests. See the Additional Results field description for more detail. (This is an LIS coded field for the EpiCenter).
Additional Results (R, 15, 1-5, 2) (EpiCenter Additional Results Only) – There are several tests that can be assigned multiple results at one time. An example of such a test is a Gram Stain test, which can be assigned results such as “Gram Negative Rod Seen” AND “Gram Positive Bacillus Seen”. These results differ from GnD test results of simply Positive or Negative. The EpiCenter maintains a user configurable list of additional test results that can apply to a test. Each additional result can also be associated with a quantity value, which precedes the additional result in the record. In the example above for other tests, the additional results and quantities reported are Many GP_BAC and Few GN_RODS. (These are LIS coded fields for the EpiCenter).
8.2.5. Comment Fields
The EpiCenter accepts and sends comment strings for patients, specimen, and isolates. Patient comments are exchanged in a comment record following a patient record, and include the patient type identifier. Specimen and isolate comments are exchanged in a comment record following the order record, and include either the specimen or isolate type identifier. An order record with both a specimen and isolate comment is shown below.
If multiple comment records are received for a patient, specimen, or isolate, then all comments of the same type are appended together and saved in the EpiCenter.
The EpiCenter can upload the description for chartable rules. This is the descriptive text for expert system rules that have been run against an isolate. The user has the ability to configure individual rules as reportable on a patient’s chart report. If a rule is configured to be chartable, the rule description will be included in an upload Comment record.
These comment records will follow after an Order record. Since the EpiCenter expert system rules only apply to isolates, these types of comments will only be part of Isolate Uploads as described in section 5.2.
The following is an example of a comment record containing a rule where <310> is the rule number and (AMX, ATM, CAZ) is a list of the affected antimicrobials.
C¦1¦ ¦<310> BDXpert Rule 310 text.( AMX, ATM, CAZ )¦E
Special Messages are messages that indicate a special condition for a BD Phoenix 100 panel. These will be uploaded in a Comment record as part of the test results for the panel. The following is an example of a comment record containing a special message with a list of the affected antimicrobials.
C¦1¦ ¦ Special message text. (AM, CAZ)¦T
Comment records not appearing after the Patient or Order records are ignored. Comments not containing the appropriate type indicator are ignored.
O¦1¦Acc123^1^MYCBTUB¦
C¦1¦ ¦This is a specimen comment ¦S C¦2¦ ¦This is an isolate comment ¦I
Comment Text (C, 4, 1, 1) – The text to be used for the patient, specimen or isolate comment. This field can be up to 1600 characters long.
Comment Type (C, 5, 1, 1) – The type of comment record. This field should be a P, S, I, E, T or R for patient, specimen, isolate, chartable rule, special message or result comment respectively.
8.2.6. Request Fields
The ASTM E_1394 defines a Request record with several fields for qualifying the requested data. Data requests can be made by specifying ranges of patient IDs, accession numbers, or sequence numbers. Data requests can also be qualified by date/time ranges.
All of these fields are supported by the BD EpiCenter interface for requests originating FROM the LIS.The field mapping table on the previous pages correctly indicates the support query fields. Examples can also be found in the Sample Messages section of this document.
Q¦1¦^Acc123¦¦MGIT_960_GND^INST_POSITIVE^MGIT_960
^2¦R¦19981019120000¦19981020120000
The sample record shown above is evaluated to request the result information for positive MGIT_960_GND tests, on instrument 2, for the specimen with an accession number of Acc123 between the dates specified. Note that all query criteria are ANDed together, so
that only records that meet ALL criteria will be uploaded in response to a query from the LIS.
The EpiCenter only requests information from the LIS on a per specimen basis. The request records generated by the EpiCenter contains only the accession number as a qualifying field as shown below. The EpiCenter request should be interpreted as requesting both patient demographics and all ordered tests for that specimen.
Q¦1¦^Acc123¦ ¦ ¦R
Starting Patient ID (Q, 3, 1, 1) – A patient ID value. This field should be filled if the data request is for either a single patient or for a range of patients. If patient ID is not to be used as a search criterion for requested data, then this field can be left blank. This field can be up to 16 characters long. In the example above, a range of patients is specified, beginning with patient ID “Pat123”.
Starting Access Number (Q, 3, 1, 2) – A specimen ID value. This field should be filled if the data request is for either a single specimen or for a range of specimen. If accession number is not to be used as a search criterion for requested data, then this field can be left blank. This field can be up to 20 characters long. In the example above, a range of specimen is specified, beginning with Accession number “Acc123”.
Starting Sequence Number (Q, 3, 1, 3) – A consumable ID value. This field should be filled if the data request is for either a single test or for a range of tests. If sequence number is not to be used as a search criterion for requested data, then this field can be left blank. This field can be up to 17 characters long. In the example above, no specific sequence numbers are requested.
Ending Patient ID (Q, 4, 1, 1) – The patient ID that ends the alphanumeric range of ID values requested. If only data for a single patient ID is requested, then this field can be left blank. This field can be up to 16 characters long. This field contains “Pat456” in the example shown above.
Ending Accession Number (Q, 4, 1, 2) – The specimen ID that ends the alphanumeric range of ID values requested. If only data for a single accession number is requested, then this field can be left blank. This field can be up to 20 characters long. This field contains “Acc456” in the example shown above.
Ending Sequence Number (Q, 4, 1, 3) – The consumable ID that ends the alphanumeric range of ID values requested. If only data for a single test is requested, then this field can be left blank. This field can be up to 17 characters long. In the example above, no specific sequence numbers are requested.
Test ID (Q, 5, 1, 1) – This is the code used to represent a particular test, and is used to request the results for this test type. This is an LIS coded field for the EpiCenter. For BD instrumented tests, the test id is the pre-defined test name listed in appendix B.
Test Status (Q, 5, 1, 2) – The status for a particular test. This field is only evaluated if the Test ID field contains valid data. This is an LIS coded field for the EpiCenter. For BD instrumented tests the status value is one of the pre-defined values listed in Appendix
B. The example record requests result information for XXX tests with a status of INST_POSITIVE.
Instrument Type (Q, 5, 1, 3) – This field indicates the type of instrumented test requested by the LIS. The EpiCenter may be connected to several instruments of different types. This field allows the LIS to request all tests run in a particular instrument type. The appropriate values for this field are listed for the Instrument Type in the result record (R, 14, 1, 1). If the value MGIT960 is used in this field, all MGIT_960_GND and all MGIT_960_AST tests are returned.
Instrument Number (Q, 5, 1, 4) – This field allows the LIS to request results for a particular instrument number. This field is only evaluated if the instrument type field is also set.
Result Qualifier (Q, 5, 1, 5) – This field allows the LIS to request results that have been modified since the last request. The only acceptable value for this field is MOD, to indicate modified results. If this field is left blank, all results that meet the remaining criteria are included. The BD EpiCenter will not accept other criteria in a query containing the MOD result qualifier.
Time Qualifier (Q, 6, 1, 1) – The use of this field is described in the ASTM E_1394 specification. The acceptable values for this field are S, R, and T which identify the content in the Starting and Ending Date/Time fields as follows:.
(S) - Specimen Collection Time
(R) – Result Time
(T) – Test Start Time
EpiCenter NOTE: If no Time Qualifier is specified, the Starting and Ending Times will, by default, refer to Result times.
In the example record times refer to result dates.
Starting Date/Time (Q, 7, 1, 1) – The starting date/time (inclusive) of the data requested. This field is formatted as specified in ASTM E_1394 specifications section
6.6.2. The date/time requested in example above is 12 pm on Oct 19, 1998.
EpiCenter NOTE: If a Starting Date/Time is specified and no Ending Date/Time is specified, data will be delivered whose Collection or Result Time is since the specified Starting Date/Time.
EpiCenter NOTE: If a Starting Date/Time specifies only the Date portion, the time will default to midnight 00:00:00.
Ending Date/Time (Q, 8, 1, 1) – The ending date/time (inclusive) of the data requested. This field is formatted as specified in ASTM E_1394 specifications section 6.6.2. The date/time requested in example above is 12 pm on Oct 20, 1998.
EpiCenter NOTE: If an Ending Date/Time specifies only the Date portion, the time will default to 23:59:59.
Request Information Status Code (Q, 13, 1, 1) – This field cancels an outstanding query. The only acceptable value for this field is ‘A’. The EpiCenter will ignore any other Request criteria fields when they appear in the same query record as a Query Cancel command.
If the string ‘ALL” appears in any Request Field, and no other Criteria are present, the EpiCenter will attempt to upload all test data in its database to the LIS.
If the string ‘ALL” appears in any Request Field, and any other Criteria are present, the EpiCenter will ignore it because an empty criteria field is an implied ‘ALL’.
8.2.7. Terminator Fields
Below is an example terminator record.
L¦1¦N
Termination Code (L, 3, 1, 1) – The termination code provides an explanation of the end of session as described in the ASTM E_1394 specification. The EpiCenter supports codes of N, Q, F and I in request responses downloaded from the LIS, as described in the ASTM specifications. However the EpiCenter and all of the BD instruments upload either N, Q or F. All other codes are ignored and treated as a normal message. Normal result uploads are indicated with N and request response uploads are indicated by F. Requests that contain an error are replied to with an empty message containing the code ‘Q’ in the terminator record. Requests that have no data satisfying the request, are responded to with a message containing only a Header and Terminator record.
The BD interface sends the results for a single test in a message. However there may be several tests that satisfy a data request downloaded by the LIS. In this instance, the BD interface uploads all tests but the last as single tests in separate messages marked as normal uploads (N). Only the last test that satisfies the request is marked as query response (F). Once the LIS receives the query response indicator (F), it has received all of the data it requested.
9. Simulator
Becton Dickinson has developed a simulator for testing the physical level LIS interface. This simulator can be configured to emulate either the LIS or instrument in the ASTM E_1381 protocol.
The simulator accepts ASTM E_1394 message strings from a file and transmit them across the serial port interface. The simulator also receives messages across the serial port interface and save the message text into a file. A full description of the simulators specifications can be found in a separate document. Becton Dickinson intends to make the simulator and supporting document available to LIS vendors to help facilitate interface development.
10. Appendix A - Sample EpiCenter Messages
10.1. LIS Order
BACTEC MGIT 960 GND tube ordered at the LIS for a new patient. Patient and test information is downloaded to the EpiCenter.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19981019184200 P¦1¦¦PatId123¦¦Doe^John^R^Jr.^Dr.¦¦19651029¦M¦¦2 Main St.^Baltimore^ MD^21211^USA¦¦(410) 316 - 4000¦JSMITH¦¦¦¦¦
¦P\AM\AMX¦¦¦¦19981015120000¦¦324¦¦¦¦¦¦¦ER¦St. Josephs Hospital
O¦1¦Acc123^ ^ ^Seq123¦¦ ^ ^ ^MGIT_960_GND¦¦
¦19981019023300¦¦¦SJB^MMF¦A¦¦¦19981019045200
¦Blood^Arm¦MJones¦(410)555–1234^(410)555–9876^(410)555–
7777¦¦¦¦¦19981020053400¦62¦¦O¦¦Nos
L¦1¦N
10.2. Host Query
The EpiCenter can receive a new accession number from a barcode on an instrumented test. The EpiCenter requests information from the LIS about that specimen.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.00¦19981019184200
Q¦1¦^Acc123¦¦¦R L¦1¦N
The LIS should respond to this request with patient demographics and ordered test information. Patient Demographics may not be included if they are not yet known. The terminator record should indicate that this is a message in response to a request.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19981019184200 P¦1¦¦PatId123¦¦Doe^John^R^Jr.^Dr.¦¦19651029¦M¦¦2 Main St.^Baltimore^ MD^21211^USA¦¦(410) 316 - 4000¦JSMITH¦¦¦¦¦
¦P\AM\AMX¦¦¦¦19981015120000¦¦324¦¦¦¦¦¦¦ER¦St. Josephs Hospital
O¦1¦Acc123^ ^ ^Seq123¦¦ ^ ^ ^MGIT_960_GND¦¦
¦19981019023300¦¦¦SJB^MMF¦A¦¦¦19981019045200
¦Blood^Arm¦MJones¦410)555–1234^410)555–9876^(410)555– 7777¦¦¦¦¦19981020053400¦62¦¦O¦¦Nos
L¦1¦F
10.3. LIS Requests Results
The LIS can request result information from the EpiCenter.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19981019184200 Q¦1¦^Acc123¦^Acc125¦MGIT_960_GND^INST_POSITIVE^MGIT960^2¦R
L¦1¦N
10.4. Results
EpiCenter will report results to the LIS in response to a request for results or asan unsolicited upload message. When results are uploaded in response to a request, the termination code in the terminator record will indicate that this is a message in response to a request.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦ V1.00 ¦19981019184200
P¦1¦¦PatId123 O¦1¦Acc123¦¦^^^MGIT_960_GND^430100065178 R¦1¦^^^GND^430100065178¦INST_POSITIVE^87¦¦¦¦¦P¦¦
¦19981019153400¦19981020145000¦MGIT960^^42^3^B/A12
L¦1¦F
10.4.1. BACTEC G&D Test results
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20060223115749 P¦1¦¦P0001¦¦Patient Name¦¦20040229¦M¦¦¦¦¦ADMIT_PHYS¦¦¦¦¦¦¦¦¦¦20060221092800¦¦¦¦¦
¦¦¦¦HOSP_SERV¦CLIENT
O¦1¦20060223001¦¦^^^PLUSAEF^449200917642¦R¦¦20060222092300¦¦
¦COLL_BY^RECD_BY¦¦¦¦20060223092300¦SAMP_TYPE^BODY¦ORD_PHYS¦¦
¦¦¦¦20060223105527
R¦1¦^^^GND^449200917642¦INST_NEGATIVE¦¦¦¦¦F¦¦¦20050201123046
¦20050208124106¦BT9000^92^32^7^A1
L¦1¦N
10.4.2. MGIT 960 G&D Test results:
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20050316152953
P¦1¦¦PatXYZ¦¦Mr. Fred A. Stare
O¦1¦ Acc456¦¦^^^MGIT_960_GND^430100065177¦R¦¦ 20041104082700¦¦¦¦¦¦¦20041104082750¦¦¦¦¦¦¦¦20050316152949
R¦1¦^^^GND_MGIT^430100065177¦INST_NEGATIVE^0¦¦¦¦¦F¦
¦¦20041029112243¦20041103110237¦MGIT960^^5^1^B/C17
L¦1¦N
10.4.3. MGIT AST Test Results
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦ V1.00 ¦19981019184200
P¦1¦¦PatId123 O¦1¦Acc123¦¦^^^MGIT_960_AST
R¦1¦^^^AST_MGIT^439400001234^P^0.5^ug/ml¦ INST_COMPLETE^105^^S¦¦¦¦¦¦¦¦19981019153400¦19981020145000¦ MGIT960^^42^3^ B/A13
R¦2¦^^^AST_MGIT^439400001234^AMX^0.5^ug/ml¦ INST_COMPLETE^142^^I¦¦¦¦¦¦¦¦19981019153400¦19981020145000¦ MGIT960^^42^3^ B/A14
R¦3¦^^^AST_MGIT^439400001234^AM^0.5^ug/ml¦ INST_COMPLETE^130^^R¦¦¦¦¦¦¦¦19981019153400¦19981020145000¦ MGIT960^^42^3^B/A15
L¦1¦N
10.4.4. Innova Results
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20110512155503
P¦1¦¦PatID1¦¦Smith Billy Bob B¦¦20110405¦M¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦ER¦Client
O¦1¦Acc055¦¦^^^INNOVA¦R¦¦20110422010000¦¦¦collecby^recby¦¦¦¦ 20110423102935¦RESP^bdysite¦phys¦¦¦¦¦¦20110512113329
R¦1¦^^^STREAK¦INST_COMPLETE¦¦¦¦¦F¦¦¦20110512113327¦ 20110512113329¦INNOVA^^^^1-1-1^^^Protocol 2_1
L¦1¦N
10.4.5. Isolate Results
Isolate results are different from test results as follows:
 The Test Id in Order field (O, 5, 1, 4) is set to “ISOLATE RESULT” instead of a real test id.
 The first result record of an isolate result is always an ID record regardless of test types.
 The ID record will contain the ID test source in field (R, 4, 1, 9).
 Unlike test uploads whch have a fixed number of results, the number of results included in an isolate can flucutuate until the isolate is finalized.
 Isolate level AST result records will contain Interpreted (R, 4, 1, 4), Expert (R, 4, 1, 5) and Final (R, 4, 1, 3) AST susceptibility values whereas test level results only contain Interpreted (R, 4, 1, 4) AST susceptibility values.
 Isolate level AST result records will contain the AST test source in field (R, 4, 1, 6).
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20060223111210
P¦1¦¦P0001¦¦Patient Name¦¦20040229¦M¦¦¦¦¦ADMIT_PHYS¦¦¦
¦¦¦¦¦¦¦20060221092800¦¦¦¦¦¦¦¦¦HOSP_SERV¦CLIENT
O¦1¦20060223001^1^ESCCOL¦¦^^^ISOLATE RESULT¦
¦¦20060222092300¦¦¦COLL_BY^RECD_BY¦¦¦
¦20060223092300¦SAMP_TYPE^BODY¦ORD_PHYS¦
¦¦¦¦¦20060223105534¦¦¦¦¦¦UNK R¦1¦^^^ID¦^ESCCOL^0000031BF0000021^^^^^^NMIC/ID-14¦¦¦¦¦F R¦2¦^^^AST^^AM¦^8^S^S^^NMIC/ID-14¦¦¦¦¦F R¦3¦^^^AST^^AMC¦^8/4^S^S^^NMIC/ID-14¦¦¦¦¦F R¦4¦^^^AST^^AN¦^<=2^S^S^^NMIC/ID-14¦¦¦¦¦F R¦5¦^^^AST^^ATM¦^<=2^S^S^^NMIC/ID-14¦¦¦¦¦F R¦6¦^^^AST^^CAZ¦^<=1^S^S^^NMIC/ID-14¦¦¦¦¦F R¦7¦^^^AST^^CIP¦^<=0.25^S^S^^NMIC/ID-14¦¦¦¦¦F R¦8¦^^^AST^^CTX¦^<=2^S^S^^NMIC/ID-14¦¦¦¦¦F R¦9¦^^^AST^^CXM¦^8^S^S^^NMIC/ID-14¦¦¦¦¦F
L¦1¦N
10.4.5.1. Isolate with Expert system changes - Blank MIC, modified SIR
The BDXpert system may change either the MIC value or the SIR value calculated on the instrument based on rules that implement generally accepted microbiology principles or suppressed reporting requirements. In the example below, 4 result values have been impacted.
Result record 2 and 14 have had the MIC removed and the SIR value set to “R” due to intrinsic resistance of this organism to the drugs. Note that the MIC field is empty. A blank value should be posted in the LIS system for this MIC.
Result record 5 has had the SIR value modified from S to X indicating that the drug result can not be reported.
Result record 17 has had the SIR value modified from S to R, indicating the drug should be reported as resistant even though it appeared susceptible in vitro.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20060223111330
P¦1¦¦P0001¦¦Patient Name¦¦20040229¦M¦¦¦¦¦ADMIT_PHYS
¦¦¦¦¦¦¦¦¦¦20060221105100¦¦Room¦¦¦¦¦¦¦HOSP_SERV¦CLIENT O¦1¦20060223002^2^STAWAR¦¦^^^ISOLATE RESULT¦¦¦ 20060222105100¦¦¦COLL_BY^RECD_BY¦¦¦¦20060223105100¦
SAMP_TYPE^BODY¦ORD_PHYS¦¦¦¦¦¦20060223105622¦¦¦¦¦¦UNK R¦1¦^^^ID¦^STAWAR^00000010002DD000^^^^^^PMIC/ID-14¦¦¦¦¦F R¦2¦^^^AST^^AM¦^^R^^R^PMIC/ID-14¦¦¦¦¦F R¦3¦^^^AST^^AMC¦^8/4^R^R^^PMIC/ID-14¦¦¦¦¦F R¦4¦^^^AST^^C¦^4^S^S^^PMIC/ID-14¦¦¦¦¦F R¦5¦^^^AST^^CC¦^<=0.25^X^S^X^PMIC/ID-14¦¦¦¦¦F R¦6¦^^^AST^^CIP¦^<=0.5^S^S^^PMIC/ID-14¦¦¦¦¦F R¦7¦^^^AST^^E¦^>4^R^R^^PMIC/ID-14¦¦¦¦¦F R¦8¦^^^AST^^FM¦^<=16^S^S^^PMIC/ID-14¦¦¦¦¦F R¦9¦^^^AST^^GM¦^<=2^S^S^^PMIC/ID-14¦¦¦¦¦F R¦10¦^^^AST^^LVX¦^<=1^S^S^^PMIC/ID-14¦¦¦¦¦F R¦11¦^^^AST^^LZD¦^2^S^S^^PMIC/ID-14¦¦¦¦¦F R¦12¦^^^AST^^MXF¦^<=1^S^S^^PMIC/ID-14¦¦¦¦¦F R¦13¦^^^AST^^OX¦^>2^R^R^^PMIC/ID-14¦¦¦¦¦F R¦14¦^^^AST^^P¦^^R^^R^PMIC/ID-14¦¦¦¦¦F R¦15¦^^^AST^^RA¦^<=0.5^S^S^^PMIC/ID-14¦¦¦¦¦F R¦16¦^^^AST^^SXT¦^<=0.5/9.5^S^S^^PMIC/ID-14¦¦¦¦¦F R¦17¦^^^AST^^VA¦^2^R^S^R^PMIC/ID-14¦¦¦¦¦F
L¦1¦N
10.4.5.2.Isolate with Resistance Markers
If an organism is identified with resistance mechanism, then the resistance mechanism code is reported as part of the identification result. Up to 5 resistance marker codes can be sent.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20060223111330
P¦1¦¦P0001¦¦Patient Name
¦¦20040229¦M¦¦¦¦¦ADMIT_PHYS¦¦¦¦¦¦¦¦¦¦20060221105100¦¦Room¦¦¦
¦¦¦¦HOSP_SERV¦CLIENT
O¦1¦20060223002^2^STAWAR¦¦^^^ISOLATE RESULT
¦¦¦20060222105100¦¦¦COLL_BY^RECD_BY¦¦¦¦20060223105100¦
SAMP_TYPE^BODY¦ORD_PHYS¦¦¦¦¦¦20060223105622¦¦¦¦¦¦UNK R¦1¦^^^ID¦^STAWAR^00000010002DD000^RM_MRSA^RM_GP_BL^^^^ PMIC/ID-14 ¦¦¦¦¦F
R¦2¦^^^AST^^AM¦^^R^^R^PMIC/ID-14¦¦¦¦¦F R¦3¦^^^AST^^AMC¦^8/4^R^R^^PMIC/ID-14¦¦¦¦¦F R¦4¦^^^AST^^C¦^4^S^S^^PMIC/ID-14¦¦¦¦¦F L¦1¦N
10.4.5.3.Isolate with Expert Comments
Certain rules in EpiCenter can produce chartable comments. These comments appear after the Order record and have a Comment Type of “E””.
There is no limit to the number of comments that could be reported.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20060223115915
P¦1¦¦P0001¦¦Patient Name
¦¦20040229¦M¦¦¦¦¦ADMIT_PHYS¦¦¦¦¦¦¦¦¦¦20060221105100¦¦Room¦¦¦
¦¦¦¦ HOSP_SERV¦CLIENT
O¦1¦20060223002^2^STAWAR¦¦^^^ISOLATE
RESULT¦¦¦20060222105100¦¦¦COLL_BY^RECD_BY¦¦¦¦20060223105100¦
SAMP_TYPE^BODY¦ORD_PHYS¦¦¦¦¦¦20060223112842¦¦¦¦¦¦UNK
C¦1¦¦<132>Staphylococcal isolates that are resistant to oxacillin (MRS) should not be treated with penicillins,
beta-lactam/beta-lactamase inhibitor combinations, cephalosporins and carbapenems.( AM, P )¦E C¦2¦¦<335>Macrolide-resistant staphylococci with an interpretation of susceptible or intermediate for clindamycin may have inducible resistance to clindamycin (MLSb resistance) or may be resistant only to macrolides (efflux-mechanism). User must perform a clindamycin induction test (disk approximation test). If positive, isolate should be reported as clindamycin resistant.
However, clindamycin may still be effective in some patients. If "D" zone is not observed, isolate should be reported as susceptible.( CC )¦E R¦1¦^^^ID¦^STAWAR^00000010002DD000^RM_MRSA^^^^^
PMIC/ID-14¦¦¦¦¦F R¦2¦^^^AST^^AM¦^^R^^R^PMIC/ID-14¦¦¦¦¦F R¦3¦^^^AST^^AMC¦^8/4^R^R^^PMIC/ID-14¦¦¦¦¦F R¦4¦^^^AST^^C¦^4^S^S^^PMIC/ID-14¦¦¦¦¦F R¦5¦^^^AST^^CC¦^<=0.25^X^S^X^PMIC/ID-14¦¦¦¦¦F R¦6¦^^^AST^^CIP¦^<=0.5^S^S^^PMIC/ID-14¦¦¦¦¦F R¦7¦^^^AST^^P¦^^R^^R^PMIC/ID-14¦¦¦¦¦F
L¦1¦N
10.4.5.4.Isolate with inferred results
The Expert system is able report results for drugs that were not actually tested on the instrument. This may be used where the drug on the test is not commonly prescribed, but is used as an indicator for drugs available in the pharmacy.
In this case, because the drug was not rested in the instrument, the inferred result records do not have MIC or test source values. In this case a blank should be posted for the MIC. Result record 6 is an inferred result in the example below.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20060223115845
P¦1¦¦P0001¦¦Patient Name
¦¦20040229¦M¦¦¦¦¦ADMIT_PHYS¦¦¦¦¦¦¦¦¦¦20060221105100¦¦Room¦¦¦
¦¦¦¦HOSP_SERV¦CLIENT
O¦1¦20060223002^1^ENTCFAA¦¦^^^ISOLATE RESULT¦¦¦20060222105100¦¦¦COLL_BY^RECD_BY¦¦¦¦20060223105100¦
SAMP_TYPE^BODY¦ORD_PHYS¦¦¦¦¦¦20060223112842¦¦¦¦¦¦UNK R¦1¦^^^ID¦^ENTCFAA^000017F82CD064C0^^^^^^PMIC/ID-14¦¦¦¦¦F R¦2¦^^^AST^^AM¦^<=2^S^S^^PMIC/ID-14¦¦¦¦¦F R¦3¦^^^AST^^C¦^8^S^S^^PMIC/ID-14¦¦¦¦¦F R¦4¦^^^AST^^E¦^2^I^I^^PMIC/ID-14¦¦¦¦¦F R¦5¦^^^AST^^GMS¦^<=500^S^S^^PMIC/ID-14¦¦¦¦¦F R¦6¦^^^AST^^IPM¦^^S^^S¦¦¦¦¦F R¦7¦^^^AST^^LVX¦^<=1^S^S^^PMIC/ID-14¦¦¦¦¦F R¦8¦^^^AST^^LZD¦^2^S^S^^PMIC/ID-14¦¦¦¦¦F R¦9¦^^^AST^^MXF¦^<=1^S^S^^PMIC/ID-14¦¦¦¦¦F
L¦1¦N
10.4.5.5.Isolate with multiple test sources
The Isolate results in EpiCenter are a consolidation of all AST results for an organism. As a result, if more than one test is used to generate these susceptibility results, more than one test source code could be reported back to the LIS system.
In the example below, the result record 16 is a drug not available on the panels tested, and was thus a supplemental test with test code “KB”.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20060223120400
P¦1¦¦P0001¦¦Patient Name
¦¦20040229¦M¦¦¦¦¦ADMIT_PHYS¦¦¦¦¦¦¦¦¦¦20060221092800¦¦¦¦¦¦¦¦¦ HOSP_SERV¦CLIENT
O¦1¦20060223003^1^KLEPNEP¦¦^^^ISOLATE RESULT¦¦¦20060222092300¦¦¦COLL_BY^RECD_BY¦¦¦¦20060223092300¦
SAMP_TYPE^BODY¦ORD_PHYS¦¦¦¦¦¦20060223120336¦¦¦¦¦¦UNK R¦1¦^^^ID¦^KLEPNEP^000012DFF9412020^^^^^^ NMIC/ID-14¦¦¦¦¦F R¦2¦^^^AST^^AM¦^>16^R^R^^NMIC/ID-14¦¦¦¦¦F R¦3¦^^^AST^^AMC¦^16/8^I^I^^NMIC/ID-14¦¦¦¦¦F R¦4¦^^^AST^^AN¦^32^I^I^^NMIC/ID-14¦¦¦¦¦F R¦5¦^^^AST^^ATM¦^<=2^S^S^^NMIC/ID-14¦¦¦¦¦F R¦6¦^^^AST^^CF¦^<=1^S^S^^NMIC-2¦¦¦¦¦F R¦7¦^^^AST^^CIP¦^2^I^I^^NMIC/ID-14¦¦¦¦¦F R¦8¦^^^AST^^CTX¦^<=2^S^S^^NMIC/ID-14¦¦¦¦¦F R¦9¦^^^AST^^CXM¦^16^I^I^^NMIC/ID-14¦¦¦¦¦F R¦10¦^^^AST^^CZ¦^16^I^I^^NMIC/ID-14¦¦¦¦¦F R¦11¦^^^AST^^FOX¦^<=1^S^S^^NMIC-2¦¦¦¦¦F R¦12¦^^^AST^^FM¦^64^I^I^^NMIC/ID-14¦¦¦¦¦F R¦13¦^^^AST^^LVX¦^2^S^S^^NMIC/ID-14¦¦¦¦¦F R¦14¦^^^AST^^MEM¦^<=1^S^S^^NMIC/ID-14¦¦¦¦¦F R¦15¦^^^AST^^NN¦^8^I^I^^NMIC-2¦¦¦¦¦F R¦16¦^^^AST^^OX¦^4^R^R^^KB¦¦¦¦¦F R¦17¦^^^AST^^PIP¦^>64^R^R^^NMIC/ID-14¦¦¦¦¦F
L¦1¦N
10.4.5.6.Isolate with multiple instruments
An EpiCenter is configured with mixed instruments, Phoenix and MGIT 960.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20050316152604 P¦1¦¦PatXYZ¦¦Mr. Fred A. Stare O¦1¦Acc456^2^STACOH¦¦^^^ISOLATE RESULT¦¦¦ 20041104082700¦¦¦¦¦¦¦ 20041104082750¦¦¦¦¦¦¦¦¦¦¦¦¦¦UNK R¦1¦^^^ID¦^STACOH^00001C71C71C71C7^^^^^^GPIDAST 954¦¦¦¦¦F R¦2¦^^^AST^^AM¦^^X^^X^GPIDAST 954¦¦¦¦¦F R¦3¦^^^AST^^AMC¦^>16/8^X^R^X^GPIDAST 954¦¦¦¦¦F R¦4¦^^^AST^^CF¦^>0.125^X^R^X^GPIDAST 954¦¦¦¦¦F R¦5¦^^^AST^^CIP¦^<=2^S^S^^GPIDAST 954¦¦¦¦¦F R¦6¦^^^AST^^CRO¦^>1^X^X^^GPIDAST 954¦¦¦¦¦F R¦7¦^^^AST^^FEP¦^<=4^X^S^X^GPIDAST 954¦¦¦¦¦F R¦8¦^^^AST^^GM¦^>32^R^R^^GPIDAST 954¦¦¦¦¦F R¦9¦^^^AST^^INH^0.10 ug/mL¦^^R^R^^MGIT_960_AST94¦¦¦¦¦F R¦10¦^^^AST^^INH^0.40 ug/mL¦^^S^S^^MGIT_960_AST94¦¦¦¦¦F R¦11¦^^^AST^^IPM¦^4^X^S^X^GPIDAST 954¦¦¦¦¦F R¦12¦^^^AST^^LVX¦^^S^^S¦¦¦¦¦F R¦13¦^^^AST^^MEM¦^<=1^X^S^X^GPIDAST 954¦¦¦¦¦F R¦14¦^^^AST^^OFX¦^^S^^S¦¦¦¦¦F R¦15¦^^^AST^^P¦^>1^X^X^^GPIDAST 954¦¦¦¦¦F R¦16¦^^^AST^^RA^1.0 ug/mL¦^^I^I^^MGIT_960_AST94¦¦¦¦¦F R¦17¦^^^AST^^SXT¦^^^^^GPIDAST 954¦¦¦¦¦F R¦18¦^^^AST^^TE¦^^^^^GPIDAST 954¦¦¦¦¦F R¦19¦^^^AST^^TZP^0.5¦^^R^R^^ASTDIA1¦¦¦¦¦F
L¦1¦N

## Table 7
11. Appendix B – Tests and Results
The following appendix lists test names and test status values. If the LIS code for a test name or test status is different than the name, it is in parenthesis next to the name. Instrumented tests and their status values appear in bold. At the EpiCenter the non-instrumented tests can be reconfigured by the user. Note: This list is not intended to be an exhaustive list, for example only some of the Test Names are listed for the Phoenix Gram Negative and Gram Positive panel types.
Test Type |  | Test Name (LIS Code) | Test Status (LIS Code)
Growth and Detection Tests |  | 30 degree MGIT 960 G&D Tube (30MGT) | False Positive
Incomplete
Negative
Positive
|  | Acid Fast Smear (AFS) | Incomplete
Acid Fast Bacilli Seen (AFB Seen)
No Acid Fast Bacilli Seen (No AFB Seen)
|  | ProbeTec ET Assay (BD_PROBETEC_GND) | Positive (INST_POSITIVE)
Negative (INST_NEGATIVE)
Indeterminate (INST_INDETERMINATE)
Equivocal (INST_EQUIVOCAL)
QC Pass (INST_QC_PASS)
QC Fail (INST_QC_FAIL)
|  | Colistin (COLIS) | Incomplete Negative Positive
Gram Stain (GRMSTN)
Lowenstein-Jensen Medium (LJ)
Myco/F Lytic (MYFLTC)
|  | MGIT 960 G&D Tube (MGIT_960_GND) | Pending (Pending)
False Positive (False Positive)
Manual Negative (Manual Negative)
Manual Positive (Manual Positive)
Threshold Positive (Threshold Positive)
Negative (INST_NEGATIVE)
Ongoing (INST_ONGOING)
Removed Ongoing (INST_REMOVED)
Positive (INST_POSITIVE)
Complete (INST_COMPLETE)
|  | Middlebrook 7H11 Medium (7H11) | Contaminate
Incomplete
Negative
Positive
ID Tests |  | Indirect Acid Fast Smear (INDAFS) | Complete Incomplete
Indirect Gram Stain (INDGRMSTN)
ProbeTec ET MAC (ETMAC)
ProbeTec ET Mtb (ETMTB)
AST – Diameter Tests |  | Ampicillin Disk (AMDISK) | Complete Incomplete
Cephalothin Disk (CFDISK)
Gentamicin Disk (GMDISK)
Penicillin Disk (PDISK)
Test Type | Test Name (LIS Code) |  |  | Test Status LIS Code
AST – MGIT 960 | MGIT 960 AST Carrier Set (MGIT_960_AST) |  |  | Pending (Pending)
Manual Complete (Manual Complete)
AST Error (INST_ERROR)
Complete (INST_COMPLETE)
Ongoing (INST_ONGOING)
AST-MIC Tests | 460TB AST (460TB) |  |  | Complete Incomplete
Ethambutol AST Test (ETHM)
Isoniazid AST Test (INH)
Pyrazinmide AST Test (PZA)
Rifampin AST Test (RIF)
Streptomycin AST Test (STREPM)
Phoenix Gram Negative | NID |  |  | Complete (INST_COMPLETE) Rapid Complete (INST_RAPID_COMPLETE) Ongoing (INST_ONGOING) Partially Complete (INST_PARTIAL_COMPLETE) In Attention Complete (INST_ATTN_COMPLETE) In Attention Ignored (INST_ATTN_IGNORED) Pending (INST_PENDING) QC Status see section 14.3.3
NQC01
NMIC/ID-1
NMIC/ID-13
NMIC/ID-14
NMIC/ID-15
Phoenix Gram Positive | PID |  |  | Complete (INST_COMPLETE) Rapid Complete (INST_RAPID_COMPLETE) Ongoing (INST_ONGOING) Partially Complete (INST_PARTIAL_COMPLETE) In Attention Complete (INST_ATTN_COMPLETE) In Attention Ignored (INST_ATTN_IGNORED) Pending (INST_PENDING) QC Status see section 14.3.3
PQC01
PMIC/ID-1
PMIC/ID-17
PMIC/ID-18
PMIC/ID-20
BD Sceptor Panels |  |  |  | Complete (COMPLETE)
BACTEC 9000 | DVE PLUS Aerobic/F (96AER) DVE PLUS Anaerobic/F (97ANA) Lytic/10 Anaerobic/F (LYTIC10ANF) Myco/F Lytic (MYCFLYTIC) Mycosis-IC/F (MYCOSISIC) PEDS Plus/F (PEDPLUSF) PLUS Aerobic/F (PLUSAEF) Plus Aerobic/F (82) (PLUS_A) Plus Aerobic/F(DVE)(80) (DVE_A) Plus Anaerobic/F (PLUSANF) Plus Anaerobic/F(83) (PLUS_N) Plus Anaerobic/F(DVE)(81) (DVE_N) Standard 10 Aerobic/F (STD10AEF) Standard Aerobic/F(90) (STDAER) Standard Anaerobic/F (STDANF) |  |  | Confirmed Positive(Confirmed Positive) Manual Negative(Manual Negative) Manual Positive(Manual Positive) Negative(INST_NEGATIVE) Ongoing (INST_ONGOING) Pending (Pending) Positive(INST_POSITIVE) Removed Ongoing(INST_REMOVED)
BACTEC FX | Lytic/10 Anaerobic/F (LYTIC10ANF) Myco/F Lytic (MYCFLYTIC) Mycosis-IC/F (MYCOSISIC) PEDS Plus/F (PEDPLUSF) PLUS Aerobic/F (PLUSAEF) Plus Anaerobic/F (PLUSANF) Standard 10 Aerobic/F (STD10AEF) |  |  | Confirmed Positive(Confirmed Positive) Manual Negative(Manual Negative) Manual Positive(Manual Positive) Negative(INST_NEGATIVE) Ongoing (INST_ONGOING) Pending (Pending) Positive(INST_POSITIVE) Removed Ongoing(INST_REMOVED)
Test Type | Test Name (LIS Code) |  |  | Test Status LIS Code
Standard Anaerobic/F (STDANF)
Innova | Innova Protocol (INNOVA) |  |  | Complete (INST_COMPLETE) Error (INST_ERROR)
Other Tests | Blood Agar Plate (TSA) |  |  | Complete Incomplete
Blood Agar Plate Anaerobic (BAPANA)
CDC Anaerobe 5% Sheep Blood Agar (CDC)
CDC Anaerobe Laked Sheep Blood Agar with KV (KV)
Cefinase (CEFINASE)
Choc Plate (CHOC)
Indole (INDOLE)
MacConkey II Agar (MAC)
12. Appendix C – Antibiotic Abbreviations
The following is a list of the antibiotics in the BD database, with the associated LIS Code. On the EpiCenter and BD Phoenix 100, these codes are configurable by the user. However on the MGIT and ProbeTec ET instruments these values are not configurable. Therefore a direct instrument interface sends these values as AST results.
1st gen cephalosporin CEFG1 Framycetin SO
1st gen cephalosporins CEF_G1 Furazolidone FX 1st gen cephalosporins w/ anti-pseudomonaCEF_H1AP Fusidic Acid FA 2nd gen cephalosporin; grp 1 GRP1 Gatifloxacin GAT 2nd gen cephalosporin; grp 2 (cephamycinsGRP2 Gemifloxacin GEM 2nd gen cephalosporin; grp 3 GRP3 Gentamicin GM 2nd gen cephalosporins CEF_G2 Gentamicin-Syn GMS
3rd gen cephalosporin w/ aerobic GNR actiCEF_MAP Glycopeptides GLYCPEP 3rd gen cephalosporin w/ anti-pseudomonalCEF_AP Grepafloxacin GRX
3rd gen cephalosporins CEF_G3 High level aminoglycosides HLAR
4th gen cephalosporin CEFG4 Hydrolyzable Penicillins PEN_GRP
4th gen cephalosporins CEF_G4 Imidazole IMIDAZ
6' methoxypenicillin MO_PEN Imipenem IPM
Amdinocillin AMD Isepamicin ISP
amdinopenicillins AMDINO_PEN Isoniazid INH
Amikacin AN Josamycin JM
Aminocyclitols AMNCYC Kanamycin K
Aminoglycosides AMNGLY Kanamycin Synergy KS
aminopenicillins AM_PEN Levofloxacin LVX aminopenicillins w/ Beta-lact. Inhib. AM_PEN_BLI Lincomycin L Amoxicillin AMX Lincosamides LINCO
Amoxicillin/Clavulanate AMC Linezolid LZD Amoxicillin/Clavulanate (f) AXC Lividomycin LV Ampicillin AM Lomefloxacin LOM
Ampicillin/Sulbactam SAM Loracarbef LOR Ampicillin/Sulbactam (f) SXA Lymecycline LYM Antimycobacterial drugs AMYCOB Macrolides MACRO
Apalcillin APL Macrolides Lincosam. StreptoMLS
Apramycin AP Mecillinam MEC
Arbekacin ARB Meropenem MEM
Aspoxicillin APX Methacycline MC
Astromycin AST Methicillin DP
Azithromycin AZM Metronidazole MET
Azlocillin AZ Mezlocillin MZ
Aztreonam ATM Mezlocillin/Sulbactam MZS
Bacitracin B Micronomycin MCR
Benzylpenicillin BZP Midecamycin MID
Beta-lactams BETA_LAC Minocycline MI
Biapenem BPM Monobactams MONOBAC
Capreomycin CAP Moxalactam MOX
Carbacephem CARBACEF Moxifloxacin MXF
Carbenicillin CB Mupirocin MUP
Carbepenems CARBAPEN Nafcillin NF
carboxypenicillins CO_PEN Nalidixic Acid NA carboxypenicillins w/ Beta-lact. Inhib. CO_PEN_BLI Neomycin N
Carumonam CAR Netilmicin NET
Cefaclor CEC Nifuroazide NFX
Cefadroxil CFR Nifurzide NZD
Cefamandole MA Nitrofurantoin FM
Cefatrizine FAT Nitroxoline NIT
Cefazolin CZ Non-hydrolyzable PenicillinsPEN_M
Cefbuperazone CFB Norfloxacin NOR
Cefdinir CDR Novobiocin NB
Cefditoren CDN Nystatin NY
Cefditoren pivoxil CDNp Ofloxacin OFX
Cefepime FEP Olandeomycin OL
Cefepime/Sulbactam SFP Ornidazole ORN
Cefetamet-pivoxil CAT Others OTHER
Cefixime CFM Oxacillin OX
Cefmenoxime CMX Oxolinic Acid OA
Cefmetazole CMZ Oxytetracycline T
Cefminox CNX p-Aminosalicylic acid PAS
Cefodizime CDZ Panipenem PAN
Cefonicid CID Paromomycin PAR
Cefoperazone CFP Pefloxacin PEF
Cefoperazone/Sulbactam SCP Penamycin PENAMYCIN
Ceforanide CND Penicillin G P
Cefotaxime CTX Penicillin V PV

## Table 8
Cefotaxime/Clavulanate | CCX | Penicillin w/ Beta-lact. inhPEN_BLI
Cefotaxime/Sulbactam | SCT | Penicillins PENS
Cefotetan | CTT | Penicillins w/ Beta-lact. inPENS_BLI
Cefotiam | CFT | Phenoxymethylpenicillin PXM
Cefotiam-hexiteil | CFTh | Pipemidic Acid PI
Cefoxitin | FOX | Piperacillin PIP
Cefozopran | CFZ | Piperacillin/Sulbactam SPI
Cefpimazole | CPZ | Piperacillin/Tazobactam TZP
Cefpiramide | CPM | Pirlimycin PRL
Cefpirome | CPO | Piromidic Acid PIR
Cefpirome/Sulbactam | SCO | Pivampicillim PAM
Cefpodoxime-proxetil | CPD | Pivmecillinam PMC
Cefpodoxime/Clavulanate | CCP | Polymyxin B PB
Cefprozil | CPR | Primary aminoglycosides PRI_AMN
Cefroxadine | CFX | Primary penicillin PRI_PEN
Cefsulodin | CFS | Primary quinolones PRI_QUIN
Ceftazadime/Sulbactam | SCZ | Pristinamycin PR
Ceftazidime | CAZ | Pyrazinamide PZA
Ceftazidime/Clavulanate | CCZ | Quinolones QUIN
Cefteram | CEM | Quinupristin/dalfopristin SYN
Ceftibuten | CTB | Rifabutin RBN
Ceftiofur | XNL | Rifampin RA
Ceftizoxime | ZOX | Rifamycin RF
Ceftriaxone | CRO | Rifamycins RIFAMYC
Ceftriaxone/Clavulanate | CCR | Rokitamycin ROK
Cefuroxime sodium | CXM | Rosoxacin R
Cefuroxime-axetil | CXMa | Roxithromycin RXT
Cefuzonam | CZN | Sarafloxacin SRF
Cephalexin | CN | Sisomycin SIS
Cephaloridine | CD | Sparfloxacin SPX
Cephalosporins | CEPHEM | Spectinomycin SPT
Cephalosporins w/ Beta-lactamase inhib. | CEPHEM_BLI | Spiramycin SP
Cephalosporins w/ Beta-lactamse inhib. | CEF_BLI | Streptogramins STREPTO
Cephalothin | CF | Streptomycin S
Cephapirin | CP | Streptomycin-Syn STS
Cephradine | CH | Sulbenicillin SBC
Chloramphenicol | C | Sulfadiazine SD
Chlortetracycline | A | Sulfamethizole TH
Cinoxacin | CIN | Sulfamethoxazole SMZ
Ciprofloxacin | CIP | sulfapenicillin SU_PEN
Clarithromycin | CLR | Sulfathiazole ST
Clinafloxacin | CLX | Sulfisoxazole G
Clindamycin | CC | Systemic fluoroquinilones SYS_FQUIN
Clofazimine | CLO | Teicoplanin TEC
Cloxacillin | CX | Telithromycin TEL
Colistin | CL | Temafloxacin TMA
Cyclic peptides | CYC_PEP | Temocillin TEM
Cycloserine | CS | Tetracycline TE
Daptomycin | DAP | Tetracyclines TET
Dibekacin | DKB | Thiacetazone TB1
Dicloxacillin | DX | Thiamphenicol TP
Difloxacin | DFX | Ticarcillin TIC
Dihydrofolate reductase inhibitor | DHFR_INH | Ticarcillin/Clavulanate TIM
Dihydrofolate reduct./Dihydropteroate synDHFR_DHPS_INH Tilmicosin TIL
Dihydropteroate synthetase inhibitor | DHPS_INH | Tobramycin NN
Dirithromycin | DTM | Tosufloxacin TFX
Doxycycline | D | Trimeth/Sulfa (DIN) STG
Enoxacin | ENX | Trimethoprim TMP
Enrofloxacin | EFX | Trimethoprim/Sulfadiazine SDT
Epicillin | EP | Trimethoprim/SulfamethoxazolSXT
Erythromycin | E | Triple Sulfa SSS
Ethambutol | EM | Trospectinomycin TRP
Ethionamide | EA | Trovafloxacin TVA
Fleroxacin | FLE | Tylosin TYL
Flomoxef | FLO | Unspecified UNS
Flucloxacillin | FO | ureidopenicillins UR_PEN
Flumequine | FLQ | ureidopenicillins w/ B-lact.UR_PEN_BLI
Fluoroquinilones | FQUIN | Urinary fluoroquinilones UR_FQUIN
Folate antagonists | FOL_ANT | Vancomycin VA
Fosfomycin w/G6PD | FF | Virginiamycin VRG

## Table 9
13. Appendix D – Organism Abbreviations
The following is a list of the organisms in the BD database, with the associated LIS Code. On the EpiCenter and BD Phoenix 100, these codes are configurable by the user. However on the MGIT and ProbeTec ET instruments these values are not configurable. Therefore a direct instrument interface sends these values as AST results.
Abiotrophia ABI | Lactobacillus catenaformis | LACBCAT
Abiotrophia defectiva STRDEF | Lactobacillus crispatus | LACBCRI
Acholeplasma ACHO | Lactobacillus delbrueckii ssp bulgaricus | LACBDELB
Acholeplasma laidlawii ACHOLAI | Lactobacillus delbrueckii ssp delbrueckii | LACBDELD
Achromobacter ACHR | Lactobacillus delbrueckii ssp lactis | LACBDELL
Achromobacter piechaudii ALCPIE | Lactobacillus fermentum | LACBFER
Achromobacter species ACHRSPE | Lactobacillus gasseri | LACBGAS
Achromobacter xylosoxidans ssp denitrificALCDEN | Lactobacillus jensenii | LACBJEN
Achromobacter xylosoxidans ssp xylosoxidaALCXYL | Lactobacillus johnsonii | LACBJOH
Acid Fast Bacilli AFB | Lactobacillus kefiri | LACBKEF
Acid Fast Bacilli (Group) AFBG | Lactobacillus paraplantarum | LACBPAR
Acidaminococcus ACIA | Lactobacillus plantarum | LACBPLA
Acidaminococcus fermentans ACIAFER | Lactobacillus reuteri | LACBREU
Acidovorax ACID | Lactobacillus rhamnosus | LACBRHA
Acidovorax delafieldii ACIDDEL | Lactobacillus salivarius | LACBSAL
Acidovorax facilis ACIDFAC | Lactobacillus species | LACBSPE
Acidovorax temperans ACIDTEM | Lactobacillus uli | LACBULI
Acinetobacter ACIN | Lactococcus | LACC
Acinetobacter baumannii ACINBAU | Lactococcus garvieae | LACCGAR
Acinetobacter baumannii/calcoaceticus comACINBCX | Lactococcus lactis ssp cremoris | LACCLACC
Acinetobacter baumannii/haemolyticus ACINBAUHAE | Lactococcus lactis ssp hordniae | LACCLACH
Acinetobacter calcoaceticus ACINCAL | Lactococcus lactis ssp lactis | LACCLACL
Acinetobacter haemolyticus ACINHAE | Lactococcus plantarum | LACCPLA
Acinetobacter johnsonii ACINJOH | Lactococcus raffinolactis | LACCRAF
Acinetobacter junii ACINJUN | Lactococcus species | LACCSPE
Acinetobacter lwoffii ACINLWO | Lautropia | LAU
Acinetobacter lwoffii/haemolyticus ACINLWOHAE | Lautropia mirabilis | LAUMIR
Acinetobacter radioresistens ACINRAD | Leclercia | LEC
Acinetobacter species ACINSPE | Leclercia adecarboxylata | LECADE
Actinobacillus ACTB | Legionella | LEG
Actinobacillus actinomycetemcomitans ACTBACT | Legionella anisa | LEGANI
Actinobacillus capsulatus ACTBCAP | Legionella birminghamensis | LEGBIR
Actinobacillus equuli ACTBEQU | Legionella bozemaniae | LEGBOZ
Actinobacillus hominis ACTBHOM | Legionella cincinnatiensis | LEGCIN
Actinobacillus lignieresii ACTBLIG | Legionella dumoffii | LEGDUM
Actinobacillus muris ACTBMUR | Legionella feeleii | LEGFEE
Actinobacillus pleuropneumoniae ACTBPLE | Legionella gormanii | LEGGOR
Actinobacillus rossii ACTBROS | Legionella hackeliae | LEGHAC
Actinobacillus seminis ACTBSEM | Legionella israelensis | LEGISR
Actinobacillus species ACTBSPE | Legionella jordanis | LEGJOR
Actinobacillus suis ACTBSUI | Legionella lansingensis | LEGLAN
Actinobacillus ureae ACTBURE | Legionella longbeachae | LEGLON
Actinobaculum ACTC | Legionella maceachernii | LEGMAC
Actinobaculum schaalii ACTCSCH | Legionella micdadei | LEGMIC
Actinobaculum suis ACTMSUI | Legionella oakridgensis | LEGOAK
Actinomadura ACTA | Legionella pneumophila | LEGPNE
Actinomadura madurae ACTAMAD | Legionella pneumophila ssp fraseri | LEGPNEF
Actinomadura pelletieri ACTAPEL | Legionella pneumophila ssp pascullei | LEGPNEPA
Actinomyces ACTM | Legionella pneumophila ssp pneumophila | LEGPNEPN
Actinomyces bovis ACTMBOV | Legionella sainthelensi | LEGSAI
Actinomyces denticolens ACTMDEN | Legionella species | LEGSPE
Actinomyces europaeus ACTMEUR | Legionella tucsonensis | LEGTUC
Actinomyces georgiae ACTMGEO | Legionella wadsworthii | LEGWAD
Actinomyces gerencseriae ACTMGER | Leifsonia | LEI
Actinomyces graevenitzii ACTMGRA | Leifsonia aquatica | CORAQU
Actinomyces hordeovulneris ACTMHOR | Leminorella | LEM
Actinomyces howelli ACTMHOW | Leminorella grimontii | LEMGRI
Actinomyces hyovaginalis ACTMHYO | Leminorella richardii | LEMRIC
Actinomyces israelii ACTMISR | Leminorella species | LEMSPE
Actinomyces meyeri ACTMMEY | Leptotrichia | LEP
Actinomyces naeslundii ACTMNAE | Leptotrichia buccalis | LEPBUC
Actinomyces neuii ssp anitratus ACTMNEUA | Leuconostoc | LEU
Actinomyces neuii ssp neuii ACTMNEUN | Leuconostoc argentinum | LEUARG
Actinomyces odontolyticus ACTMODO | Leuconostoc carnosum | LEUCAR
Actinomyces radingae ACTMRAD | Leuconostoc citreum | LEUCIT
Actinomyces slackii ACTMSLA | Leuconostoc gelidum | LEUGEL
Actinomyces species ACTMSPE | Leuconostoc lactis | LEULAC
Actinomyces turicensis ACTMTUR | Leuconostoc mesenteroides | LEUMES
Actinomyces viscosus ACTMVIS | Leuconostoc mesenteroides ssp cremoris | LEUMESC
Aerococcus AERC | Leuconostoc mesenteroides ssp dextranicum | LEUMESD
Aerococcus species AERCSPE | Leuconostoc mesenteroides ssp mesenteroides | LEUMESM
Aerococcus urinae AERCURI | Leuconostoc pseudomesenteroides | LEUPSE
Aerococcus viridans AERCVIR | Leuconostoc species | LEUSPE
Aeromonas AERM | Listeria | LIS
Aeromonas allosaccharophila AERMALL | Listeria grayi | LISGRA
Aeromonas caviae AERMCAV | Listeria innocua | LISINN
Aeromonas hydrophila AERMHYD | Listeria ivanovii | LISIVA
Aeromonas hydrophila group AERMHYDGR | Listeria ivanovii ssp ivanovii | LISIVAI
Aeromonas jandaei AERMJAN | Listeria ivanovii ssp londoniensis | LISIVAL
Aeromonas media AERMMED | Listeria monocytogenes | LISMON
Aeromonas salmonicida AERMSAL | Listeria monocytogenes/innocua | LISMONINN
Aeromonas salmonicida ssp achromogenes AERMSALA | Listeria murrayi | LISMUR
Aeromonas salmonicida ssp masoucida AERMSALM | Listeria murrayi/grayi | LISMURGRA
Aeromonas salmonicida ssp salmonicida AERMSALSA | Listeria seeligeri | LISSEE
Aeromonas salmonicida ssp smithia AERMSALSM | Listeria species | LISSPE
Aeromonas schubertii AERMSCH | Listeria welshimeri | LISWEL
Aeromonas sobria AERMVERS | Listonella | LISO
Aeromonas species AERMSPE | Listonella anguillarum | LISOANG
Aeromonas trota AERMTRO | Listonella pelagia | LISOPEL
Aeromonas veronii AERMVERV | Macrococcus | MAC
Afipia AFI | Macrococcus caseolyticus | STACAS

## Table 10
Afipia broomeae | AFIBRO | Malassezia | MAL
Afipia clevelandensis | AFICLE | Malassezia furfur | MALFUR
Afipia felis | AFIFEL | Malassezia globosa | MALGLO
Agrobacterium | AGR | Malassezia obtusa | MALOBT
Agrobacterium radiobacter | AGRRAD | Malassezia pachydermatis | MALPAC
Agrobacterium tumefaciens | AGRTUM | Malassezia restricta | MALRES
Alcaligenes | ALC | Malassezia sloofiae | MALSLO
Alcaligenes faecalis | ALCFAE | Malassezia species | MALSPE
Alcaligenes faecalis ssp faecalis | ALCFAEF | Malassezia sympodialis | MALSYM
Alcaligenes faecalis type II | ALCFAEII | Mannheimia | MAN
Alcaligenes species | ALCSPE | Mannheimia haemolytica | PASHAE
Alloiococcus | ALL | Megamonas | MEGM
Alloiococcus otitidis | ALLOTI | Megamonas hypermegale | MEGMHYP
Anaerobiospirillum | ANAB | Megasphaera | MEGS
Anaerobiospirillum succiniciproducens | ANABSUC | Megasphaera elsdenii | MEGSELS
Anaerobiospirillum thomasii | ANABTHO | Methylobacterium | MET
Anaerorhabdus | ANAR | Methylobacterium extorquens | METEXT
Anaerorhabdus furcosa | ANARFUR | Methylobacterium fujisawaense | METFUJ
Arcanobacterium | ARCA | Methylobacterium mesophilicum | METMES
Arcanobacterium bernardiae | ARCABER | Methylobacterium species | METSPE
Arcanobacterium haemolyticum | ARCAHAE | Microbacterium | MICB
Arcanobacterium pyogenes | ACTMPYO | Microbacterium arborescens | MICBARB
Arcobacter | ARCO | Microbacterium imperiale | MICBIMP
Arcobacter butzleri | ARCOBUT | Micrococcus | MIC
Arcobacter cryaerophilus | ARCOCRY | Micrococcus luteus | MICLUT
Arcobacter nitrofigilis | ARCONIT | Micrococcus lylae | MICLYL
Arcobacter skirrowii | ARCOSKI | Micrococcus species | MICSPE
Arthrobacter | ART | Micromonas | MICM
Arthrobacter agilis | ARTAGI | Micromonas micros | PEPSMIC
Arthrobacter creatinolyticus | ARTCRE | Mitsuokella | MIT
Arthrobacter cumminsii | ARTCUM | Mitsuokella multiacida | MITMUL
Arthrobacter woluwensis | ARTWOL | Mobiluncus | MOB
Aspergillus | ASP | Mobiluncus curtisii | MOBCUR
Aspergillus fumigatus | ASPFUM | Mobiluncus mulieris | MOBMUL
Aspergillus species | ASPSPE | Mobiluncus species | MOBSPE
Atopobium | ATO | Moellerella | MOE
Atopobium minutum | ATOMIN | Moellerella wisconsensis | MOEWIS
Atopobium parvulum | ATOPAR | Mogibacterium | MOG
Atopobium rimae | LACBRIM | Mogibacterium timidum | EUBTIM
Aureobacterium | AURC | Moraxella | MORA
Aureobacterium resistens | AURCRES | Moraxella (Branhamella) catarrhalis | MORABRACAT
Aureobasidium | AUR | Moraxella atlantae | MORAATL
Aureobasidium pullulans | AURPUL | Moraxella bovis | MORABOV
Bacillus | BACI | Moraxella canis | MORACAN
Bacillus amyloliquefaciens | BACIAMY | Moraxella caviae | MORACAV
Bacillus anthracis | BACIANT | Moraxella cuniculi | MORACUN
Bacillus cereus | BACICER | Moraxella lacunata | MORALAC
Bacillus circulans | BACICIR | Moraxella lincolnii | MORALIN
Bacillus coagulans | BACICOA | Moraxella nonliquefaciens | MORANON
Bacillus firmus | BACIFIR | Moraxella osloensis | MORAOSL
Bacillus licheniformis | BACILIC | Moraxella ovis | MORAOVI
Bacillus megaterium | BACIMEG | Moraxella species | MORASPE
Bacillus mycoides | BACIMYC | Morganella | MORG
Bacillus pasteurii | BACIPAS | Morganella morganii | MORGMOR
Bacillus pumilus | BACIPUM | Morganella morganii ssp morganii | MORGMORM
Bacillus species | BACISPE | Morganella morganii ssp morganii biogroup 1 | MORGMORM1
Bacillus sphaericus | BACISPH | Morganella morganii ssp sibonii | MORGMORS
Bacillus stearothermophilus | BACISTE | Mycobacterium | MYCB
Bacillus subtilis | BACISUB | Mycobacterium abscessus | MYCBABS
Bacillus subtilis ssp spizizenii | BACISUBSP | Mycobacterium africanum | MYCBAFR
Bacillus subtilis ssp subtilis | BACISUBSU | Mycobacterium asiaticum | MYCBASI
Bacillus thuringiensis | BACITHU | Mycobacterium aurum | MYCBAUR
Bacteroides | BACT | Mycobacterium avium | MYCBAVI
Bacteroides caccae | BACTCAC | Mycobacterium avium complex | MYCBMAC
Bacteroides capillosus | BACTCAP | Mycobacterium avium ssp avium | MYCBAVIA
Bacteroides coagulans | BACTCOA | Mycobacterium avium ssp paratuberculosis | MYCBAVIP
Bacteroides distasonis | BACTDIS | Mycobacterium avium ssp silvaticum | MYCBAVIS
Bacteroides distasonis group | BACTDISGR | Mycobacterium bovis | MYCBBOV
Bacteroides eggerthii | BACTEGG | Mycobacterium branderi | MYCBBRA
Bacteroides forsythus | BACTFOR | Mycobacterium celatum | MYCBCEL
Bacteroides fragilis | BACTFRA | Mycobacterium chelonae | MYCBCHE
Bacteroides fragilis group | BACTFRAGR | Mycobacterium confluentis | MYCBCON
Bacteroides merdae | BACTMER | Mycobacterium farcinogenes | MYCBFAR
Bacteroides ovatus | BACTOVA | Mycobacterium flavescens | MYCBFLA
Bacteroides putredinis | BACTPUT | Mycobacterium fortuitum | MYCBFOR
Bacteroides splanchnicus | BACTSPL | Mycobacterium fortuitum complex | MYCBFORC
Bacteroides stercoris | BACTSTI | Mycobacterium gadium | MYCBGAD
Bacteroides tectus | BACTTEC | Mycobacterium gastri | MYCBGAS
Bacteroides thetaiotaomicron | BACTTHE | Mycobacterium genavense | MYCBGEN
Bacteroides uniformis | BACTUNI | Mycobacterium gordonae | MYCBGOR
Bacteroides ureolyticus | BACTURE | Mycobacterium haemophilum | MYCBHAE
Bacteroides vulgatus | BACTVUL | Mycobacterium heidelbergense | MYCBHEI
Balneatrix | BAL | Mycobacterium interjectum | MYCBINJ
Balneatrix alpica | BALALP | Mycobacterium intermedium | MYCBINM
Bartonella | BAR | Mycobacterium intracellulare | MYCBINC
Bartonella bacilliformis | BARBAC | Mycobacterium kansasii | MYCBKAN
Bartonella clarridgeiae | BARCLA | Mycobacterium leprae | MYCBLEP
Bartonella elizabethae | BARELI | Mycobacterium malmoense | MYCBMAL
Bartonella henselae | BARHEN | Mycobacterium marinum | MYCBMAR
Bartonella quintana | BARQUI | Mycobacterium microti | MYCBMIC
Bartonella species | BARSPE | Mycobacterium mucogenicum | MYCBMUC
Bergeyella | BER | Mycobacterium neoaurum | MYCBNEO
Bergeyella zoohelcum | BERZOO | Mycobacterium nonchromogenicum | MYCBNON
Bifidobacterium | BIF | Mycobacterium obuense | MYCBOBU
Bifidobacterium adolescentis | BIFADO | Mycobacterium other than tuberculosis | MYCBMOTT
Bifidobacterium angulatum | BIFANG | Mycobacterium peregrinum | MYCBPER
Bifidobacterium bifidum | BIFBIF | Mycobacterium phlei | MYCBPHL
Bifidobacterium breve | BIFBRE | Mycobacterium rhodesiae | MYCBRHO
Bifidobacterium catenulatum | BIFCAT | Mycobacterium scrofulaceum | MYCBSCR
Bifidobacterium denticolens | BIFDEC | Mycobacterium shimoidei | MYCBSHI
Bifidobacterium dentium | BIFDEN | Mycobacterium simiae | MYCBSIM
Bifidobacterium gallicum | BIFGAL | Mycobacterium smegmatis | MYCBSME
Bifidobacterium infantis | BIFINF | Mycobacterium species | MYCBSPE
Bifidobacterium inopinatum | BIFINO | Mycobacterium szulgai | MYCBSZU
Bifidobacterium longum | BIFLON | Mycobacterium terrae | MYCBTER
Bifidobacterium pseudocatenulatum | BIFPSC | Mycobacterium terrae complex | MYCBTERC
Bifidobacterium pseudolongum ssp globosumBIFPSEG Mycobacterium thermoresistibile MYCBTHE

## Table 11
Bifidobacterium pseudolongum ssp pseudoloBIFPSEP | Mycobacterium triviale | MYCBTRI
Bifidobacterium species BIFSPE | Mycobacterium tuberculosis | MYCBTUB
Bilophila BIL | Mycobacterium tuberculosis complex | MYCBTUBC
Bilophila wadsworthia BILWAD | Mycobacterium ulcerans | MYCBULC
Blastoschizomyces BLA | Mycobacterium vaccae | MYCBVAC
Blastoschizomyces capitatus BLACAP | Mycobacterium xenopi | MYCBXEN
Bordetella BOR | Mycoplasma | MYCP
Bordetella avium BORAVI | Mycoplasma (Group) | MYCOPG
Bordetella bronchiseptica BORBROS | Mycoplasma (Group2) | MYCOP
Bordetella hinzii BORHIN | Mycoplasma buccale | MYCPBUC
Bordetella holmesii BORHOL | Mycoplasma faucium | MYCPFAU
Bordetella parapertussis BORPAR | Mycoplasma fermentans | MYCPFER
Bordetella pertussis BORPER | Mycoplasma gallisepticum | MYCPGAL
Bordetella trematum BORTRE | Mycoplasma genitalium | MYCPGEN
Borrelia BOI | Mycoplasma hominis | MYCPHOM
Borrelia species BOISPE | Mycoplasma lipophilum | MYCPLIP
Brachyspira BRAC | Mycoplasma orale | MYCPORA
Brachyspira aalborgi BRACAAL | Mycoplasma penetrans | MYCPPEN
Brachyspira hyodysenteriae BRACHYO | Mycoplasma pirum | MYCPPIR
Brachyspira innocens BRACINN | Mycoplasma pneumoniae | MYCPPNE
Brachyspira pilosicoli BRACPIL | Mycoplasma primatum | MYCPPRI
Bradyrhizobium BRAD | Mycoplasma salivarium | MYCPSAL
Bradyrhizobium japonicum BRADJAP | Mycoplasma spermatophilum | MYCPSPR
Brevibacillus BRES | Mycoplasma synoviae | MYCPSYN
Brevibacillus brevis BACIBRE | Myroides | MYR
Brevibacillus laterosporus BACILAT | Myroides odoratimimus | MYRODI
Brevibacterium BREI | Myroides odoratus | MYRODA
Brevibacterium casei BREICAS | Myroides odoratus/odoratimimus | MYRODAODI
Brevibacterium epidermidis BREIEPI | Neisseria | NEI
Brevibacterium linens BREILIN | Neisseria canis | NEICAN
Brevibacterium mcbrellneri BREIMCB | Neisseria cinerea | NEICIN
Brevibacterium species BREISPE | Neisseria elongata | NEIELO
Brevundimonas BREU | Neisseria elongata ssp elongata | NEIELOE
Brevundimonas diminuta BREUDIM | Neisseria elongata ssp glycolytica | NEIELOG
Brevundimonas vesicularis BREUVES | Neisseria elongata ssp nitroreducens | NEIELON
Brucella BRU | Neisseria flavescens | NEIFLA
Brucella abortus BRUABO | Neisseria gonorrhoeae | NEIGON
Brucella canis BRUCAN | Neisseria iguanae | NEIIGU
Brucella melitensis BRUMEL | Neisseria lactamica | NEILAC
Brucella neotomae BRUNEO | Neisseria meningitidis | NEIMEN
Brucella ovis BRUOVI | Neisseria mucosa | NEIMUC
Brucella species BRUSPE | Neisseria polysaccharea | NEIPOL
Brucella suis BRUSUI | Neisseria sicca | NEISIC
Budvicia BUD | Neisseria species | NEISPE
Budvicia aquatica BUDAQU | Neisseria subflava | NEISUB
Burkholderia BUR | Neisseria subflava biovar flava | NEISUBF
Burkholderia caryophylli BURCAR | Neisseria subflava biovar perflava | NEISUBP
Burkholderia cepacia BURCEP | Neisseria subflava biovar subflava | NEISUBS
Burkholderia cepacia (CF) BURCEPCF | Neisseria weaveri | NEIWEA
Burkholderia cepacia/Ralstonia pickettii BURCEPRALPIC | Nesterenkonia | NES
Burkholderia gladioli BURGLA | Nesterenkonia halobia | NESHAL
Burkholderia glathei BURGLT | Nocardia | NOCA
Burkholderia graminis BURGRA | Nocardia asteroides | NOCAAST
Burkholderia mallei BURMAL | Nocardia brasiliensis | NOCABRA
Burkholderia multivorans BURMUL | Nocardia brevicatena | NOCABRE
Burkholderia phenazinium BURPHE | Nocardia carnea | NOCACAR
Burkholderia pseudomallei BURPSE | Nocardia farcinica | NOCAFAR
Burkholderia pyrrocinia BURPYR | Nocardia nova | NOCANOV
Burkholderia species/Ralstonia species BURSPERALSPE | Nocardia otitidiscaviarum | NOCAOTI
Burkholderia/Ralstonia BUR/RAL | Nocardia pseudobrasiliensis | NOCAPSE
Buttiauxella BUTT | Nocardia seriolae | NOCASER
Buttiauxella agrestis BUTTAGR | Nocardia species | NOCASPE
Buttiauxella ferragutiae BUTTFER | Nocardia transvalensis | NOCATRA
Buttiauxella gaviniae BUTTGAV | Nocardia vaccinii | NOCAVAC
Butyrivibrio BUTY | Nocardiopsis | NOCO
Butyrivibrio crossotus BUTYCRO | Nocardiopsis dassonvillei | NOCODAS
Butyrivibrio fibrisolvens BUTYFIB | Obesumbacterium | OBE
Calymmatobacterium CALY | Obesumbacterium proteus | OBEPRO
Campylobacter CAM | Ochrobactrum | OCH
Campylobacter coli CAMCOL | Ochrobactrum anthropi | OCHANT
Campylobacter concisus CAMCON | Oenococcus | OEN
Campylobacter curvus CAMCUR | Oenococcus oeni | OENOEN
Campylobacter fetus CAMFET | Oerskovia | OER
Campylobacter fetus ssp fetus CAMFETF | Oerskovia species | OERSPE
Campylobacter fetus ssp venerealis CAMFETV | Oerskovia xanthineolytica | OERXAN
Campylobacter gracilis CAMGRA | Oligella | OLI
Campylobacter helveticus CAMHEL | Oligella species | OLISPE
Campylobacter hyointestinalis ssp hyointeCAMHYOH | Oligella ureolytica | OLIURO
Campylobacter hyointestinalis ssp lawsoniCAMHYOL | Oligella urethralis | OLIURT
Campylobacter jejuni CAMJEJ | Ornithobacterium | ORN
Campylobacter jejuni ssp doylei CAMJEJD | Ornithobacterium rhinotracheale | ORNRHI
Campylobacter jejuni ssp jejuni CAMJEJJ | Paenibacillus | PAE
Campylobacter lari CAMLAR | Paenibacillus alvei | PAEALV
Campylobacter mucosalis CAMMUC | Paenibacillus durus | PAEDUR
Campylobacter rectus CAMRET | Paenibacillus macerans | PAEMAC
Campylobacter rectus/curvus CAMRETCUR | Paenibacillus polymyxa | BACIPOL
Campylobacter showae CAMSHO | Pantoea | PAN
Campylobacter species CAMSPE | Pantoea agglomerans | PANAGG
Campylobacter sputorum ssp bubulus CAMSPUB | Pantoea ananatis | ERWANA
Campylobacter sputorum ssp fecalis CAMSPUF | Pantoea dispersa | PANDIS
Campylobacter sputorum ssp paraureolyticuCAMSPUP | Pantoea stewartii ssp indologenes | PANSTEI
Campylobacter sputorum ssp sputorum CAMSPUS | Pantoea stewartii ssp stewartii | PANSTES
Campylobacter upsaliensis CAMUPS | Pasteurella | PAS
Candida CAN | Pasteurella aerogenes | PASAER
Candida aaseri CANAAS | Pasteurella bettyae | PASBET
Candida albicans CANALB | Pasteurella caballi | PASCAB
Candida boidinii CANBOI | Pasteurella canis | PASCAN
Candida catenulata CANCAT | Pasteurella dagmatis | PASDAG
Candida ciferrii CANCIF | Pasteurella gallinarum | PASGAL
Candida colliculosa CANCOL | Pasteurella multocida | PASMUL
Candida dubliniensis CANDUB | Pasteurella multocida ssp gallicida | PASMULG
Candida famata CANFAM | Pasteurella multocida ssp multocida | PASMULM
Candida glabrata TORGLA | Pasteurella multocida ssp septica | PASMULS
Candida guilliermondii CANGUI | Pasteurella pneumotropica | PASPNE
Candida haemulonii CANHAE | Pasteurella species | PASSPE
Candida holmii CANHOL | Pasteurella stomatis | PASSTO
Candida inconspicua CANINC | Pasteurella trehalosi | PASTRE

## Table 12
Candida kefyr | CANKEF | Pasteurella volantium | PASVOL
Candida krusei | CANKRU | Pediococcus | PED
Candida lambica | CANLAM | Pediococcus acidilactici | PEDACI
Candida lipolytica | CANLIP | Pediococcus damnosus | PEDDAM
Candida lusitaniae | CANLUS | Pediococcus dextrinicus | PEDDEX
Candida magnoliae | CANMAG | Pediococcus parvulus | PEDPAR
Candida melibiosica | CANMEL | Pediococcus pentosaceus | PEDPEN
Candida norvegensis | CANNOR | Pediococcus species | PEDSPE
Candida parakrusei | CANPARK | Peptococcus | PEPC
Candida parapsilosis | CANPARP | Peptococcus niger | PEPCNIG
Candida pelliculosa | CANPEL | Peptostreptococcus | PEPS
Candida pintolopesii | CANPIN | Peptostreptococcus anaerobius | PEPSANA
Candida pulcherrima | CANPUL | Peptostreptococcus asaccharolyticus | PEPSASA
Candida rugosa | CANRUG | Peptostreptococcus barnesae | PEPSBAR
Candida species | CANSPE | Peptostreptococcus harei | PEPSHAR
Candida sphaerica | CANSPH | Peptostreptococcus hydrogenalis | PEPSHYD
Candida stellatoidea | CANSTE | Peptostreptococcus indolicus | PEPSIND
Candida tropicalis | CANTRO | Peptostreptococcus ivorii | PEPSIVO
Candida utilis | CANUTI | Peptostreptococcus lacrimalis | PEPSLAR
Candida valida | CANVAL | Peptostreptococcus lactolyticus | PEPSLAT
Candida viswanathii | CANVIS | Peptostreptococcus octavius | PEPSOCT
Candida zeylanoides | CANZEY | Peptostreptococcus prevotii | PEPSPRE
Capnocytophaga | CAP | Peptostreptococcus species | PEPSSPE
Capnocytophaga canimorsus | CAPCAN | Peptostreptococcus tetradius | PEPSTET
Capnocytophaga cynodegmi | CAPCYN | Peptostreptococcus vaginalis | PEPSVAG
Capnocytophaga gingivalis | CAPGIN | Photobacterium | PHOB
Capnocytophaga granulosa | CAPGRA | Photobacterium damselae | PHOBDAM
Capnocytophaga haemolytica | CAPHAE | Photobacterium damselae ssp damselae | PHOBDAMD
Capnocytophaga ochracea | CAPOCH | Photobacterium damselae ssp piscicida | PHOBDAMP
Capnocytophaga species | CAPSPE | Photorhabdus | PHOR
Capnocytophaga sputigena | CAPSPU | Photorhabdus luminescens | PHORLUM
Cardiobacterium | CAR | Pichia | PIC
Cardiobacterium hominis | CARHOM | Pichia angusta | PICANG
Catonella | CATO | Pichia anomala | HANANO
Catonella morbi | CATOMOR | Pichia fermentans | PICFER
CDC | CDC | Plesiomonas | PLE
CDC group DF-3 | CDCDF3 | Plesiomonas shigelloides | PLESHI
CDC group EF-4 | CDCEF4 | Porphyromonas | POR
CDC group EF-4a | CDCEF4a | Porphyromonas asaccharolytica | PORASA
CDC group EF-4b | CDCEF4b | Porphyromonas canoris | PORCAN
CDC group EO-2 | CDCEO2 | Porphyromonas catoniae | PORCAT
CDC group EO-3 | CDCEO3 | Porphyromonas circumdentaria | PORCIR
CDC group II b | CDCIIb | Porphyromonas endodontalis | POREND
CDC group II g | CDCIIg | Porphyromonas gingivalis | PORGIN
CDC group II h | CDCIIh | Porphyromonas levii | PORLEV
CDC group II i | CDCIIi | Porphyromonas macacae | PORMAC
CDC group IV | CDCIV | Pragia | PRA
CDC group NO-1 | CDCNO1 | Pragia fontium | PRAFON
CDC group O-1 | CDCO1 | Prevotella | PRE
CDC group O-2 | CDCO2 | Prevotella albensis | PREALB
CDC group OFBA-1 | CDCOFBA1 | Prevotella bivia | PREBIV
CDC group Vb-3 | CDCVb3 | Prevotella brevis | PREBRE
CDC group WO-1 | CDCWO1 | Prevotella bryantii | PREBRY
Cedecea | CED | Prevotella buccae | PREBUE
Cedecea davisae | CEDDAV | Prevotella buccalis | PREBUL
Cedecea lapagei | CEDLAP | Prevotella corporis | PRECOR
Cedecea neteri | CEDNET | Prevotella dentalis | PREDEA
Cedecea species | CEDSPE | Prevotella denticola | PREDEI
Cedecea species 3 | CEDSPE3 | Prevotella disiens | PREDIS
Cedecea species 5 | CEDSPE5 | Prevotella enoeca | PREENO
Cellulomonas | CEL | Prevotella heparinolytica | PREHEP
Cellulomonas humilata | ACTMHUM | Prevotella intermedia | PREINT
Cellulomonas turbata | OERTUR | Prevotella loescheii | PRELOE
Centipeda | CEN | Prevotella melaninogenica | PREMEL
Centipeda periodontii | CENPER | Prevotella nigrescens | PRENIG
Chromobacterium | CHRO | Prevotella oralis | PREORA
Chromobacterium violaceum | CHROVIO | Prevotella oris | PREORI
Chryseobacterium | CHRB | Prevotella oulorum | PREOUL
Chryseobacterium gleum | CHRBGLE | Prevotella ruminicola | PRERUM
Chryseobacterium indologenes | CHRBIND | Prevotella tannerae | PRETAN
Chryseobacterium meningosepticum | CHRBMEN | Prevotella veroralis | PREVER
Chryseobacterium scophthalmum | CHRBSCO | Prevotella zoogleoformans | PREZOO
Chryseomonas | CHRM | Propionibacterium | PROB
Citrobacter | CIT | Propionibacterium acidipropionici | PROBACI
Citrobacter amalonaticus | CITAMA | Propionibacterium acnes | PROBACN
Citrobacter braakii | CITBRA | Propionibacterium avidum | PROBAVI
Citrobacter farmeri | CITFAR | Propionibacterium avidum ssp granulosum | PROBAVIG
Citrobacter freundii | CITFRE | Propionibacterium freundenreichii ssp freundenrPROBFREF
Citrobacter freundii complex | CITFRECX | Propionibacterium freundenreichii ssp shermaniiPROBFRES
Citrobacter gillenii | CITSPE10 | Propionibacterium granulosum | PROBGRA
Citrobacter koseri | CITKOS | Propionibacterium jensenii | PROBJEN
Citrobacter murliniae | CITSPE11 | Propionibacterium lymphophilum | PROBLYM
Citrobacter rodentium | CITSPE9 | Propionibacterium propionicum | PROBPRO
Citrobacter sedlakii | CITSED | Propionibacterium species | PROBSPE
Citrobacter species | CITSPE | Propioniferax | PROF
Citrobacter werkmanii | CITWER | Propioniferax innocua | PROFINN
Citrobacter youngae | CITYOU | Proteus | PROT
Clostridium | CLO | Proteus mirabilis | PROTMIR
Clostridium acetobutylicum | CLOACO | Proteus myxofaciens | PROTMYX
Clostridium amniovalericum | CLOAMN | Proteus penneri | PROTPEN
Clostridium argentinense | CLOARG | Proteus species | PROTSPE
Clostridium aurantibutyricum | CLOAUR | Proteus vulgaris | PROTVUL
Clostridium baratii | CLOBAR | Proteus vulgaris/penneri | PROTVULPEN
Clostridium beijerinckii | CLOBEI | Prototheca | PROH
Clostridium bifermentans | CLOBIF | Prototheca wickerhamii | PROHWIC
Clostridium botulinum | CLOBOT | Prototheca zopfii | PROHZOP
Clostridium butyricum | CLOBUT | Providencia | PROV
Clostridium cadaveris | CLOCAD | Providencia alcalifaciens | PROVALC
Clostridium carnis | CLOCAR | Providencia heimbachae | PROVHEI
Clostridium celatum | CLOCEL | Providencia rettgeri | PROVRET
Clostridium cellobioparum | CLOCEO | Providencia rustigianii | PROVRUS
Clostridium chauvoei | CLOCHA | Providencia species | PROVSPE
Clostridium clostridioforme | CLOCLO | Providencia stuartii | PROVSTU
Clostridium cochlearium | CLOCOH | Providencia stuartii urea+ | PROVSTUUR
Clostridium cocleatum | CLOCOL | Pseudomonas | PSE
Clostridium difficile | CLODIF | Pseudomonas aeruginosa | PSEAER
Clostridium fallax | CLOFAL | Pseudomonas aeruginosa (CF) | PSEAERCF
Clostridium felsineum CLOFEL Pseudomonas alcaligenes PSEALC
Clostridium ghonii CLOGHO Pseudomonas fluorescens PSEFLU Clostridium glycolicum CLOGLY Pseudomonas fluorescens/putida PSEFLUPUT Clostridium haemolyticum CLOHAE Pseudomonas luteola CHRMLUT
Clostridium hastiforme CLOHAS Pseudomonas mendocina PSEMEN
Clostridium histolyticum CLOHIS Pseudomonas monteilii PSEMON Clostridium indolis CLOIND Pseudomonas oryzihabitans FLAIORY Clostridium innocuum CLOINN Pseudomonas pertucinogena PSEPER Clostridium irregulare CLOIRR Pseudomonas pseudoalcaligenes PSEPSE Clostridium leptum CLOLEP Pseudomonas pseudoalcaligenes ssp pseudoalcaligPSEPSEP
Clostridium limosum CLOLIM Pseudomonas putida PSEPUT
Clostridium malenominatum CLOMAL Pseudomonas species PSESPE
Clostridium mangenotii CLOMAN Pseudomonas stutzeri PSESTU
Clostridium nexile CLONEX Pseudonocardia PSEN
Clostridium novyi CLONOV Pseudonocardia autotrophica PSENAUT
Clostridium novyi A CLONOVA Pseudoramibacter PSER
Clostridium novyi B CLONOVB Pseudoramibacter alactolyticus EUBALA Clostridium oceanicum CLOOCE Psychrobacter PSY
Clostridium orbiscindens CLOORB Psychrobacter immobilis PSYIMM Clostridium oroticum CLOORO Psychrobacter phenylpyruvicus MORAPHE Clostridium paraputrificum CLOPAR Rahnella RAH
Clostridium perfringens CLOPER Rahnella aquatilis RAHAQU
Clostridium putrefaciens CLOPUR Ralstonia RAL
Clostridium putrificum CLOPUT Ralstonia paucula CDCIVC2
Clostridium ramosum CLORAM Ralstonia pickettii BURPIC Clostridium sardiniense CLOSAR Ralstonia pickettii biovar 1 BURPIC1 Clostridium sartagoforme CLOSAT Ralstonia pickettii biovar 2 BURPIC2 Clostridium scatologenes CLOSCA Ralstonia pickettii biovar 3 BURPIC3 Clostridium septicum CLOSEP Ralstonia solanacearum BURSOL
Clostridium sordellii CLOSOR Rhodococcus RHOC
Clostridium species CLOSPE Rhodococcus coprophilus RHOCCOP
Clostridium sphenoides CLOSPH Rhodococcus equi RHOCEQU
Clostridium spiroforme CLOSPI Rhodococcus erythropolis RHOCERY
Clostridium sporogenes CLOSPO Rhodococcus fascians RHOCFAS
Clostridium sporosphaeroides CLOSPS Rhodococcus globerulus RHOCGLO
Clostridium subterminale CLOSUB Rhodococcus rhodochrous RHOCRHO
Clostridium symbiosum CLOSYM Rhodococcus species RHOCSPE
Clostridium tertium CLOTER Rhodotorula RHOT
Clostridium tetani CLOTET Rhodotorula glutinis RHOTGLU Clostridium tyrobutyricum CLOTYR Rhodotorula mucilaginosa var mucilaginosa RHOTMUCM Collinsella COLL Rhodotorula rubra RHOTRUB
Collinsella aerofaciens EUBAER Rhodotorula species RHOTSPE
Comamonas COM Rickettsia RICK
Comamonas terrigena COMTER Rickettsia (Group) RICKG
Comamonas testosteroni COMTES Riemerella RIE
Coprococcus COP Riemerella anatipestifer RIEANA
Coprococcus catus COPCAT Rikenella RIK
Coprococcus comes COPCOM Rikenella microfusus RIKMIC
Coprococcus eutactus COPEUT Roseomonas ROS
Corynebacterium COR Roseomonas cervicalis ROSCER
Corynebacterium accolens CORACC Roseomonas fauriae ROSFAU
Corynebacterium afermentans CORAFE Roseomonas gilardii ROSGIL
Corynebacterium afermentans ssp afermentaCORAFEA Roseomonas species ROSSPE
Corynebacterium afermentans ssp lipophiluCORAFEL Rothia ROT
Corynebacterium amycolatum CORAMY Rothia dentocariosa ROTDEN
Corynebacterium amycolatum/minitissimum CORAMYMIN Rothia mucilaginosa STOMUC
Corynebacterium amycolatum/striatum CORAMYSTR Ruminococcus RUM
Corynebacterium argentoratense CORARG Ruminococcus albus RUMALB
Corynebacterium auris CORAUR Ruminococcus bromii RUMBRO
Corynebacterium bovis CORBOV Ruminococcus hansenii STRHAN
Corynebacterium coyleae CORCOY Ruminococcus productus RUMPRO
Corynebacterium cystitidis CORCYS Saccharomyces SAC
Corynebacterium diphtheriae CORDIP Saccharomyces cerevisiae SACCER
Corynebacterium diphtheriae ssp belfanti CORDIPB Saccharomyces pastorianus SACPAS
Corynebacterium diphtheriae ssp gravis CORDIPG Saccharomyces uvarum SACUVA
Corynebacterium diphtheriae ssp intermediCORDIPI Salmonella SAL
Corynebacterium diphtheriae ssp mitis CORDIPM Salmonella aberdeen SALABE
Corynebacterium durum CORDUR Salmonella abortus-equi SALABOE
Corynebacterium flavescens CORFLA Salmonella adelaide SALADL
Corynebacterium genitalium CORGEN Salmonella aderike SALADR
Corynebacterium genitalium biovar I CORGENI Salmonella agona SALAGO
Corynebacterium genitalium biovar II CORGENII Salmonella alachua SALALA
Corynebacterium genitalium biovar III CORGENIII Salmonella anatum SALANA
Corynebacterium genitalium biovar IV CORGENIV Salmonella arizonae SALARI
Corynebacterium genitalium biovar V CORGENV Salmonella avana SALAVA
Corynebacterium genitalium biovar VI CORGENVI Salmonella bahrenfeld SALBAH
Corynebacterium glucuronolyticum CORGLC Salmonella blockley SALBLO
Corynebacterium imitans CORIMI Salmonella bongori SALBON
Corynebacterium jeikeium CORJEI Salmonella braenderup SALBRA
Corynebacterium kutscheri CORKUT Salmonella bredeney SALBRE
Corynebacterium macginleyi CORMAC Salmonella bunn SALBUN
Corynebacterium matruchotii CORMAT Salmonella california SALCAL
Corynebacterium minutissimum CORMIN Salmonella carrau SALCAR
Corynebacterium mucifaciens CORMUC Salmonella cerro SALCER
Corynebacterium mycetoides CORMYC Salmonella champaign SALCHA
Corynebacterium pilosum CORPIL Salmonella chittagong SALCHI
Corynebacterium propinquum CORPRO Salmonella choleraesuis SALCHO Corynebacterium pseudodiphtheriticum CORPSD Salmonella choleraesuis ssp arizonae SALCHOA Corynebacterium pseudogenitalium CORPSG Salmonella choleraesuis ssp choleraesuis SALCHOC Corynebacterium pseudogenitalium biovar CCORPSGC1 Salmonella choleraesuis ssp diarizonae SALCHOD Corynebacterium pseudogenitalium biovar CCORPSGC2 Salmonella choleraesuis ssp houtenae SALCHOH Corynebacterium pseudogenitalium biovar CCORPSGC3 Salmonella choleraesuis ssp indica SALCHOI Corynebacterium pseudogenitalium biovar CCORPSGC4 Salmonella choleraesuis ssp salamae SALCHOS Corynebacterium pseudotuberculosis CORPST Salmonella cubana SALCUB
Corynebacterium renale CORREN Salmonella dakar SALDAK
Corynebacterium renale group CORRENGR Salmonella daressalaam SALDAR
Corynebacterium riegelii CORRIE Salmonella derby SALDER
Corynebacterium seminale CORSEM Salmonella dessau SALDES
Corynebacterium singulare CORSIN Salmonella dublin SALDUB
Corynebacterium species CORSPE Salmonella duesseldorf SALDUE
Corynebacterium striatum CORSTR Salmonella enteritidis SALENT
Corynebacterium thomssenii CORTHO Salmonella fresno SALFRE
Corynebacterium ulcerans CORULC Salmonella gallinarum SALGAL
Corynebacterium urealyticum CORURE Salmonella give SALGIV
Corynebacterium variabile CORVAR Salmonella haardt SALHAA
Corynebacterium vitaeruminis CORVIT Salmonella hadar SALHAD
Corynebacterium xerosis CORXER Salmonella hamburg SALHAM

## Table 13
Coxiella | COX | Salmonella heidelberg | SALHEI
Coxiella burnetii | COXBUR | Salmonella illinois | SALILL
Cryptococcus | CRY | Salmonella infantis | SALINF
Cryptococcus albidus | CRYALB | Salmonella inverness | SALINV
Cryptococcus albidus var aerius | CRYALBAE | Salmonella java | SALJAA
Cryptococcus albidus var albidus | CRYALBAL | Salmonella javiana | SALJAI
Cryptococcus albidus var diffluens | CRYALBD | Salmonella kirkee | SALKIR
Cryptococcus gastricus | CRYGAS | Salmonella kunduchi | SALKUN
Cryptococcus laurentii | CRYLAU | Salmonella kvittingfoss | SALKVI
Cryptococcus luteolus | CRYLUT | Salmonella lansing | SALLAN
Cryptococcus neoformans | CRYNEO | Salmonella litchfield | SALLIT
Cryptococcus neoformans var gattii | CRYNEOG | Salmonella liverpool | SALLIV
Cryptococcus neoformans var neoformans | CRYNEON | Salmonella london | SALLON
Cryptococcus species | CRYSPE | Salmonella luciana | SALLUC
Cryptococcus terreus | CRYTER | Salmonella manhattan | SALMAN
Cryptococcus uniguttulatus | CRYUNI | Salmonella meleagridis | SALMEL
Deinococcus | DEI | Salmonella memphis | SALMEM
Deinococcus radiodurans | DEIRAD | Salmonella michigan | SALMIC
Delftia | DELF | Salmonella minneapolis | SALMIA
Delftia acidovorans | COMACI | Salmonella minnesota | SALMIS
Dermabacter | DERB | Salmonella montevideo | SALMON
Dermabacter hominis | DERBHOM | Salmonella muenchen | SALMUE
Dermacoccus | DERC | Salmonella newington | SALNEI
Dermacoccus nishinomiyaensis | MICNIS | Salmonella newport | SALNEP
Dermatophilus | DERT | Salmonella nottingham | SALNOT
Dermatophilus congolensis | DERTCON | Salmonella ohio | SALOHI
Desulfomonas | DESM | Salmonella onderstepoort | SALOND
Desulfomonas pigra | DESMPIG | Salmonella oranienburg | SALORA
Desulfovibrio | DESV | Salmonella paratyphi A | SALPARA
Desulfovibrio species | DESVSPE | Salmonella paratyphi B | SALPARB
Dialister | DIA | Salmonella poona | SALPOO
Dialister pneumosintes | BACTPNE | Salmonella pullorum | SALPUL
Dichelobacter | DIC | Salmonella quinhon | SALQUI
Dichelobacter nodosus | DICNOD | Salmonella rubislaw | SALRUB
Dietzia | DIE | Salmonella saintpaul | SALSAI
Dietzia maris | DIEMAR | Salmonella schwarzengrund | SALSCH
DO NOT USE | STRINS | Salmonella senftenberg | SALSEN
Dolosigranulum | DOL | Salmonella species | SALSPE
Dolosigranulum pigrum | DOLPIG | Salmonella tallahassee | SALTAL
Edwardsiella | EDW | Salmonella thompson | SALTHO
Edwardsiella hoshinae | EDWHOS | Salmonella typhi | SALTYP
Edwardsiella ictaluri | EDWICT | Salmonella typhimurium | SALTYPM
Edwardsiella tarda | EDWTAR | Salmonella virginia | SALVIR
Edwardsiella tarda biogroup 1 | EDWTAR1 | Salmonella westerstede | SALWES
Eggerthella | EGG | Salmonella worthington | SALWOR
Eggerthella lenta | EUBLEN | Sarcina | SAR
Ehrlichia | EHR | Sarcina species | SARSPE
Ehrlichia chaffeensis | EHRCHA | Sarcina ventriculi | SARVEN
Ehrlichia equi | EHREQU | Sebaldella | SEB
Ehrlichia sennetsu | EHRSEN | Sebaldella termitidis | SEBTER
Eikenella | EIK | Selenomonas | SEL
Eikenella corrodens | EIKCOR | Selenomonas artemidis | SELART
Empedobacter | EMP | Selenomonas dianae | SELDIA
Empedobacter brevis | EMPBRE | Selenomonas flueggei | SELFLU
Enterobacter | ENTB | Selenomonas infelix | SELINF
Enterobacter aerogenes | ENTBAER | Selenomonas noxia | SELNOX
Enterobacter aerogenes/cloacae | ENTBAERCLO | Selenomonas sputigena | SELSPU
Enterobacter amnigenus | ENTBAMN | Serratia | SER
Enterobacter amnigenus biogroup 1 | ENTBAMN1 | Serratia entomophila | SERENT
Enterobacter amnigenus biogroup 2 | ENTBAMN2 | Serratia ficaria | SERFIC
Enterobacter asburiae | ENTBASB | Serratia fonticola | SERFON
Enterobacter cancerogenus | ENTBCAN | Serratia grimesii | SERGRI
Enterobacter cloacae | ENTBCLO | Serratia liquefaciens | SERLIQ
Enterobacter dissolvens | ENTBDIS | Serratia marcescens | SERMAR
Enterobacter gergoviae | ENTBGER | Serratia odorifera | SERODO
Enterobacter hormaechei | ENTBHOR | Serratia odorifera 1 | SERODO1
Enterobacter intermedius | ENTBINT | Serratia odorifera 2 | SERODO2
Enterobacter kobei | ENTBKOB | Serratia plymuthica | SERPLY
Enterobacter nimipressuralis | ENTBNIM | Serratia proteamaculans ssp proteamaculans | SERPROP
Enterobacter sakazakii | ENTBSAK | Serratia proteamaculans ssp quinovora | SERPROQ
Enterobacter species | ENTBSPE | Serratia rubidaea | SERRUB
Enterococcus | ENTC | Serratia species | SERSPE
Enterococcus avium | ENTCAVI | Shewanella | SHE
Enterococcus avium HRE | ENTCAVIHR | Shewanella algae | SHEALG
Enterococcus casseliflavus | ENTCCAS | Shewanella putrefaciens | SHEPUT
Enterococcus casseliflavus HRE | ENTCCASHR | Shewanella putrefaciens biovar 1 | SHEPUT1
Enterococcus casseliflavus Van C | ENTCCASVC | Shewanella putrefaciens biovar 2 | SHEPUT2
Enterococcus casseliflavus/gallinarum | ENTCCASGAL | Shewanella putrefaciens biovar 3 | SHEPUT3
Enterococcus cecorum | ENTCCEC | Shigella | SHI
Enterococcus columbae | ENTCCOL | Shigella boydii | SHIBOY
Enterococcus dispar | ENTCDIS | Shigella dysenteriae | SHIDYS
Enterococcus durans | ENTCDUR | Shigella flexneri | SHIFLE
Enterococcus durans HRE | ENTCDURHR | Shigella sonnei | SHISON
Enterococcus durans/faecium | ENTCDURFAI | Shigella species | SHISPE
Enterococcus faecalis | ENTCFAA | Simonsiella | SIM
Enterococcus faecalis HRE | ENTCFAAHR | Simonsiella muelleri | SIMMUE
Enterococcus faecalis Van A | ENTCFAAVA | Skermania | SKE
Enterococcus faecalis Van B | ENTCFAAVB | Skermania piniformis | SKEPIN
Enterococcus faecalis VRE | ENTCFAAVR | Slackia | SLA
Enterococcus faecium | ENTCFAI | Slackia exigua | EUBEXI
Enterococcus faecium HRE | ENTCFAIHR | Slackia heliotrinireducens | PEPSHEL
Enterococcus faecium Van A | ENTCFAIVA | Sphingobacterium | SPHB
Enterococcus faecium Van B | ENTCFAIVB | Sphingobacterium multivorum | SPHBMUL
Enterococcus faecium VRE | ENTCFAIVR | Sphingobacterium multivorum/thalpophilum | SPHBMULTHA
Enterococcus flavescens | ENTCFLA | Sphingobacterium species | SPHBSPE
Enterococcus gallinarum | ENTCGAL | Sphingobacterium spiritivorum | SPHBSPI
Enterococcus gallinarum Van C | ENTCGALVC | Sphingobacterium thalpophilum | SPHBTHA
Enterococcus hirae | ENTCHIR | Sphingomonas | SPHM
Enterococcus hirae/faecium | ENTCHIRFAI | Sphingomonas capsulata | SPHMCAP
Enterococcus malodoratus | ENTCMAL | Sphingomonas paucimobilis | SPHMPAU
Enterococcus mundtii | ENTCMUN | Spirochete | SPRC
Enterococcus pseudoavium | ENTCPSE | Spirochete (Group) | SPRCG
Enterococcus raffinosus | ENTCRAF | Sporobolomyces | SPOB
Enterococcus raffinosus Van A | ENTCRAFVA | Sporobolomyces salmonicolor | SPOBSAL
Enterococcus saccharolyticus | ENTCSAC | Staphylococcus | STA
Enterococcus solitarius | ENTCSOL | Staphylococcus arlettae | STAARL
Enterococcus species | ENTCSPE | Staphylococcus aureus | STAAUE
Enterococcus sulfureus | ENTCSUL | Staphylococcus aureus MR |  | STAAUEMR
Erwinia | ERW | Staphylococcus aureus ssp anaerobius |  | STAAUEAN
Erwinia persicina | ERWPER | Staphylococcus aureus ssp aureus |  | STAAUEAU
Erysipelothrix | ERY | Staphylococcus aureus Van I |  | STAAUEVI
Erysipelothrix rhusiopathiae | ERYRHU | Staphylococcus aureus Van R |  | STAAUEVR
Erysipelothrix tonsillarum | ERYTON | Staphylococcus auricularis |  | STAAUI
Escherichia | ESC | Staphylococcus capitis |  | STACAI
Escherichia blattae | ESCBLA | Staphylococcus capitis ssp capitis |  | STACAIC
Escherichia coli | ESCCOL | Staphylococcus capitis ssp ureolyticus |  | STACAIU
Escherichia coli AD | ESCCOLAD | Staphylococcus caprae |  | STACAP
Escherichia coli atypical | ESCCOLATY | Staphylococcus carnosus |  | STACAR
Escherichia coli H2S+ | ESCCOLHS | Staphylococcus carnosus ssp carnosus |  | STACARC
Escherichia coli inactive | ESCCOLINA | Staphylococcus carnosus ssp utilis |  | STACARU
Escherichia coli serotype O111 | ESCCOL0111 | Staphylococcus chromogenes |  | STACHR
Escherichia coli serotype O157 | ESCCOL0157 | Staphylococcus chromogenes/hyicus |  | STACHRHYI
Escherichia fergusonii | ESCFER | Staphylococcus coag negative |  | STACNEG
Escherichia hermannii | ESCHER | Staphylococcus cohnii |  | STACOH
Escherichia vulneris | ESCVUL | Staphylococcus cohnii ssp cohnii |  | STACOHC
Eubacterium | EUB | Staphylococcus cohnii ssp urealyticum |  | STACOHU
Eubacterium barkeri | EUBBAR | Staphylococcus delphini |  | STADEL
Eubacterium biforme | EUBBIF | Staphylococcus epidermidis |  | STAEPI
Eubacterium brachy | EUBBRA | Staphylococcus epidermidis MR |  | STAEPIMR
Eubacterium combesii | EUBCOM | Staphylococcus equorum |  | STAEQU
Eubacterium contortum | EUBCON | Staphylococcus felis |  | STAFEL
Eubacterium cylindroides | EUBCYL | Staphylococcus gallinarum |  | STAGAL
Eubacterium dolichum | EUBDOL | Staphylococcus haemolyticus |  | STAHAE
Eubacterium eligens | EUBELI | Staphylococcus haemolyticus MR |  | STAHAEMR
Eubacterium formicigenerans | EUBFOR | Staphylococcus hominis |  | STAHOM
Eubacterium hadrum | EUBHAD | Staphylococcus hominis MR |  | STAHOMMR
Eubacterium hallii | EUBHAL | Staphylococcus hominis ssp hominis |  | STAHOMH
Eubacterium limosum | EUBLIM | Staphylococcus hominis ssp novobiosepticus |  | STAHOMN
Eubacterium minutum | EUBMIN | Staphylococcus hyicus |  | STAHYI
Eubacterium moniliforme | EUBMON | Staphylococcus intermedius |  | STAINT
Eubacterium multiforme | EUBMUL | Staphylococcus kloosii |  | STAKLO
Eubacterium nitritogenes | EUBNIT | Staphylococcus lentus |  | STALEN
Eubacterium nodatum | EUBNOD | Staphylococcus lugdunensis |  | STALUG
Eubacterium ramulus | EUBRAM | Staphylococcus lutrae |  | STALUT
Eubacterium rectale | EUBREC | Staphylococcus muscae |  | STAMUS
Eubacterium saburreum | EUBSAB | Staphylococcus pasteuri |  | STAPAS
Eubacterium saphenum | EUBSAP | Staphylococcus piscifermentans |  | STAPIS
Eubacterium siraeum | EUBSIR | Staphylococcus pulvereri |  | STAPUL
Eubacterium species | EUBSPE | Staphylococcus saccharolyticus |  | STASAC
Eubacterium sulci | FUSSUL | Staphylococcus saprophyticus |  | STASAP
Eubacterium tenue | EUBTEN | Staphylococcus saprophyticus MR |  | STASAPMR
Eubacterium tortuosum | EUBTOR | Staphylococcus saprophyticus ssp bovis |  | STASAPB
Eubacterium ventriosum | EUBVEN | Staphylococcus saprophyticus ssp saprophyticus |  | STASAPS
Eubacterium yurii ssp margaretiae | EUBYURM | Staphylococcus schleiferi |  | STASCH
Eubacterium yurii ssp schtitka | EUBYURS | Staphylococcus schleiferi ssp coagulans |  | STASCHC
Eubacterium yurii ssp yurii | EUBYURY | Staphylococcus schleiferi ssp schleiferi |  | STASCHS
Ewingella | EWI | Staphylococcus sciuri |  | STASCI
Ewingella americana | EWIAME | Staphylococcus sciuri ssp carnaticus |  | STASCIC
Exiguobacterium | EXI | Staphylococcus sciuri ssp rodentium |  | STASCIR
Exiguobacterium acetylicum | EXIACE | Staphylococcus sciuri ssp sciuri |  | STASCIS
Facklamia | FAC | Staphylococcus simulans |  | STASIM
Facklamia hominis | FACHOM | Staphylococcus simulans MR |  | STASIMMR
Fibrobacter | FIB | Staphylococcus species |  | STASPE
Fibrobacter succinogenes | FIBSUC | Staphylococcus species MR |  | STASPEMR
Filifactor | FILF | Staphylococcus vitulinus |  | STAVIT
Filifactor alocis | FUSALO | Staphylococcus warneri |  | STAWAR
Filobasidiella | FIL | Staphylococcus warneri MR |  | STAWARMR
Filobasidiella neoformans | FILNEO | Staphylococcus xylosus |  | STAXYL
Finegoldia | FIN | Stenotrophomonas |  | STE
Finegoldia magna | PEPSMAG | Stenotrophomonas africana |  | STEAFR
Flavimonas | FLAI | Stenotrophomonas maltophilia |  | STEMAL
Flavobacterium | FLAO | Stenotrophomonas maltophilia (CF) |  | STEMALCF
Flavobacterium mizutaii | FLAOMIZ | Stomatococcus |  | STO
Flavobacterium species | FLAOSPE | Stomatococcus species |  | STOSPE
Francisella | FRA | Streptobacillus |  | STRB
Francisella novicida | FRANOV | Streptobacillus moniliformis |  | STRBMON
Francisella philomiragia | FRAPHI | Streptococcus |  | STR
Francisella tularensis | FRATUL | Streptococcus acidominimus |  | STRACI
Francisella tularensis ssp holarctica | FRATULH | Streptococcus agalactiae |  | STRAGA
Francisella tularensis ssp tularensis | FRATULT | Streptococcus alactolyticus |  | STRALA
Fungi | FUNGI | Streptococcus anginosus |  | STRANG
Fungi (Group) | FUNGIG | Streptococcus bovis |  | STRBOV
Fusobacterium | FUS | Streptococcus bovis I |  | STRBOVI
Fusobacterium gonidiaformans | FUSGON | Streptococcus bovis II |  | STRBOVII
Fusobacterium mortiferum | FUSMOR | Streptococcus canis |  | STRCAN
Fusobacterium naviforme | FUSNAV | Streptococcus constellatus |  | STRCON
Fusobacterium necrogenes | FUSNEG | Streptococcus criceti |  | STRCRC
Fusobacterium necrophorum | FUSNEC | Streptococcus cristatus |  | STRCRS
Fusobacterium necrophorum ssp funduliformFUSNECF |  | Streptococcus downei |  | STRDOW
Fusobacterium necrophorum ssp necrophorumFUSNECN |  | Streptococcus dysgalactiae |  | STRDYS
Fusobacterium nucleatum FUSNUC |  | Streptococcus dysgalactiae ssp dysgalactiae |  | STRDYSDY
Fusobacterium nucleatum ssp animalis FUSNUCA |  | Streptococcus dysgalactiae ssp equisimilis |  | STRDYSEM
Fusobacterium nucleatum ssp fusiforme FUSNUCF |  | Streptococcus equi |  | STREQU
Fusobacterium nucleatum ssp nucleatum FUSNUCN |  | Streptococcus equi ssp equi |  | STREQUE
Fusobacterium nucleatum ssp polymorphum FUSNUCP |  | Streptococcus equi ssp zooepidemicus |  | STREQUZ
Fusobacterium nucleatum ssp vincentii FUSNUCV |  | Streptococcus equinus |  | STREQN
Fusobacterium perfoetens FUSPEF |  | Streptococcus ferus |  | STRFER
Fusobacterium periodonticum FUSPER |  | Streptococcus gordonii |  | STRGOR
Fusobacterium prausnitzii FUSPRA |  | Streptococcus group C |  | STRGRC
Fusobacterium russii FUSRUS |  | Streptococcus group C/G |  | STRGRCG
Fusobacterium ulcerans FUSULC |  | Streptococcus group CFG |  | STRGRCFG
Fusobacterium varium FUSVAR |  | Streptococcus group E |  | STRGRE
Gardnerella GAR |  | Streptococcus group F |  | STRGRF
Gardnerella vaginalis GARVAG |  | Streptococcus group G |  | STRGRG
Gemella GEM |  | Streptococcus group L |  | STRGRL
Gemella bergeri GEMBER |  | Streptococcus hyointestinalis |  | STRHYO
Gemella haemolysans GEMHAE |  | Streptococcus iniae |  | STRINI
Gemella morbillorum GEMMOR |  | Streptococcus intermedius |  | STRINR
Gemella sanguinis GEMSAN |  | Streptococcus macedonicus |  | STRMAC
Gemella species GEMSPE |  | Streptococcus milleri group |  | STRMILGR
Geotrichum GEO |  | Streptococcus mitis |  | STRMIT
Geotrichum candidum GEOCAN |  | Streptococcus mitis group |  | STRMITGR
Geotrichum species GEOSPE |  | Streptococcus mitis/pneumoniae |  | STRMITPNE
Globicatella GLO |  | Streptococcus mutans |  | STRMUT

## Table 14
Globicatella sanguinis GLOSAN | Streptococcus mutans group | STRMUTGR
Gordonia GOR | Streptococcus oralis | STRORA
Gordonia aichiensis GORAIC | Streptococcus parasanguinis | STRPAR
Gordonia amarae GORAMA | Streptococcus pleomorphus | STRPLE
Gordonia bronchialis GORBRO | Streptococcus pneumoniae | STRPNE
Gordonia rubropertincta GORRUB | Streptococcus pneumoniae Pen I | STRPNEPI
Gordonia sputi GORSPU | Streptococcus pneumoniae Pen R | STRPNEPR
Gordonia terrae GORTER | Streptococcus porcinus | STRPOR
Gram + Bacilli G+B | Streptococcus pyogenes | STRPYO
Gram + Bacteria + | Streptococcus ratti | STRRAT
Gram + Cocci G+C | Streptococcus salivarius | STRSAL
Gram + Coccobacilli G+CB | Streptococcus salivarius group | STRSALGR
Gram - Bacilli G-B | Streptococcus sanguinis | STRSAN
Gram - Bacteria - | Streptococcus sanguinis group | STRSANGR
Gram - Cocci G-C | Streptococcus sobrinus | STRSOB
Gram - Coccobacilli G-CB | Streptococcus species | STRSPE
Gram - Diplococci G-DC | Streptococcus suis | STRSUI
Granulicatella GRAN | Streptococcus thermophilus | STRTHE
Granulicatella adiacens STRADJ | Streptococcus uberis | STRUBE
Haemophilus HAE | Streptococcus vestibularis | STRVES
Haemophilus aphrophilus HAEAPH | Streptococcus viridans group | STRVIRGR
Haemophilus aphrophilus/paraphrophilus HAEAPHPAR | Succinivibrio | SUC
Haemophilus ducreyi HAEDUC | Succinivibrio dextrinosolvens | SUCDEX
Haemophilus haemoglobinophilus HAEHAG | Sutterella | SUTE
Haemophilus haemolyticus HAEHAL | Sutterella wadsworthensis | SUTEWAD
Haemophilus influenzae HAEINF | Suttonella | SUT
Haemophilus influenzae biogroup aegyptiusHAEINFA | Suttonella indologenes | SUTIND
Haemophilus influenzae biotype I HAEINFI | Tatumella | TAT
Haemophilus influenzae biotype II HAEINFII | Tatumella ptyseos | TATPTY
Haemophilus influenzae biotype III HAEINFIII | Tetragenococcus | TET
Haemophilus influenzae biotype IV HAEINFIV | Tetragenococcus halophilus | TETHAL
Haemophilus influenzae biotype V HAEINFV | Tissierella | TIS
Haemophilus influenzae biotype VI HAEINFVI | Tissierella praeacuta | TISPRA
Haemophilus influenzae biotype VIII HAEINFVIII | Torulopsis | TOR
Haemophilus parahaemolyticus HAEPAH | Torulopsis candida | TORCAN
Haemophilus parainfluenzae HAEPAI | Torulopsis pintolopesii | TORPIN
Haemophilus parainfluenzae biotype I HAEPAII | Trabulsiella | TRA
Haemophilus parainfluenzae biotype II HAEPAIII | Trabulsiella guamensis | TRAGUA
Haemophilus parainfluenzae biotype III HAEPAIIII | Trichosporon | TRI
Haemophilus parainfluenzae biotype IV HAEPAIIV | Trichosporon aquatile | TRIAQU
Haemophilus paraphrophilus HAEPAP | Trichosporon beigelii | TRIBEI
Haemophilus parasuis HAEPAS | Trichosporon pullulans | TRIPUL
Haemophilus segnis HAESEG | Tsukamurella | TSU
Haemophilus somnus HAESOM | Tsukamurella inchonensis | TSUINC
Haemophilus species HAESPE | Tsukamurella paurometabola | TSUPAU
Hafnia HAF | Tsukamurella pulmonis | TSUPUL
Hafnia alvei HAFALV | Tsukamurella tyroinosolvens | TSUTYR
Hafnia alvei group 1 HAFALV1 | Tsukamurella wratislaviensis | TSUWRA
Hallella HAL | Turicella | TUR
Hallella seregens HALSER | Turicella otitidis | TUROTI
Halomonas HALO | Unidentified organism | UNIORG
Halomonas aquamarina ALCFAEH | Unspecified | UNS
Hansenula HAN | Ureaplasma | URE
Hansenula polymorpha HANPOL | Ureaplasma urealyticum | UREURE
Hansenula saturnas HANSAT | Vagococcus | VAG
Helcococcus HELC | Vagococcus species | VAGSPE
Helcococcus kunzii HELCKUN | Veillonella | VEI
Helicobacter HELI | Veillonella atypica | VEIATY
Helicobacter acinonychis HELIACI | Veillonella dispar | VEIDIS
Helicobacter cinaedi HELICIN | Veillonella parvula | VEIPAR
Helicobacter felis HELIFEL | Veillonella species | VEISPE
Helicobacter fennelliae HELIFEN | Vibrio | VIB
Helicobacter muridarum HELIMUR | Vibrio alginolyticus | VIBALG
Helicobacter mustelae HELIMUS | Vibrio cholerae | VIBCHO
Helicobacter nemestrinae HELINEM | Vibrio cincinnatiensis | VIBCIN
Helicobacter pullorum HELIPUL | Vibrio fluvialis | VIBFLU
Helicobacter pylori HELIPYL | Vibrio furnissii | VIBFUR
Histoplasma HIS | Vibrio harveyi | VIBHAR
Histoplasma capsulatum HISCAP | Vibrio hollisae | VIBHOL
Issatchenkia ISS | Vibrio metschnikovii | VIBMET
Issatchenkia orientalis ISSORI | Vibrio mimicus | VIBMIM
Johnsonella JOH | Vibrio parahaemolyticus | VIBPAR
Johnsonella ignava JOHIGN | Vibrio species | VIBSPE
Jonesia JON | Vibrio vulnificus | VIBVUL
Jonesia denitrificans JONDEN | Weeksella | WEE
Kingella KIN | Weeksella virosa | WEEVIR
Kingella denitrificans KINDEN | Weeksella virosa/Bergeyella zoohelcum | WEEVIRBERZOO
Kingella kingae KINKIN | Weeksella/Bergeyella | WEE/BER
Kingella oralis KINORA | Weissella | WEI
Kingella species KINSPE | Weissella confusa | WEICON
Klebsiella KLE | Weissella paramesenteroides | WEIPAR
Klebsiella granulomatis CALYGRA | Wolinella | WOLI
Klebsiella ornithinolytica KLEORN | Wolinella succinogenes | WOLISUC
Klebsiella oxytoca KLEOXY | Xenorhabdus | ZEN
Klebsiella oxytoca/pneumoniae KLEOXYPNE | Xenorhabdus beddingii | XENBED
Klebsiella planticola KLEPLA | Xenorhabdus bovienii | XENBOV
Klebsiella pneumoniae ssp ozaenae KLEPNEO | Xenorhabdus nematophilus | XENNEM
Klebsiella pneumoniae ssp pneumoniae KLEPNEP | Xenorhabdus poinarii | XENPOI
Klebsiella pneumoniae ssp rhinoscleromatiKLEPNER | Yamadazyma | YAM
Klebsiella species KLESPE | Yamadazyma guilliermondii | YAMGUI
Klebsiella terrigena KLETER | Yarrowia | YAR
Kloeckera KLO | Yarrowia lipolytica | YARLIP
Kloeckera apiculata KLOAPI | Yeast | YST
Kluyvera KLU | Yeast (Group) | YSTG
Kluyvera ascorbata KLUASC | Yeast-like | YSTL
Kluyvera cryocrescens KLUCRY | Yeast-like (Group) | YSTLG
Kluyvera georgiana KLUGEO | Yersinia | YER
Kluyvera species KLUSPE | Yersinia aldovae | YERALD
Kluyveromyces KLUM | Yersinia bercovieri | YERBER
Kluyveromyces lactis KLUMLAC | Yersinia enterocolitica | YERENT
Kluyveromyces marxianus KLUMMAR | Yersinia enterocolitica group | YERENTGR
Kocuria KOC | Yersinia frederiksenii | YERFRE
Kocuria kristinae MICKRI | Yersinia intermedia | YERINT
Kocuria rosea MICROS | Yersinia kristensenii | YERKRI
Kocuria varians MICVAR | Yersinia mollaretii | YERMOL
Kurthia KUR | Yersinia pestis | YERPES
Kurthia gibsonii KURGIB | Yersinia pseudotuberculosis | YERPSE

## Table 15
Kurthia zopfii | KURZOP | Yersinia rohdei | YERROH
Kytococcus | KYT | Yersinia ruckeri | YERRUC
Kytococcus sedentarius | MICSED | Yersinia species | YERSPE
Lactobacillus | LACB | Yokenella | YOK
Lactobacillus acidophilus | LACBACI | Yokenella regensburgei | YOKREG
Lactobacillus brevis | LACBBRE | Zygosaccharomyces | ZYG
Lactobacillus buchneri | LACBBUC | Zygosaccharomyces bailli | ZYGBAI
Lactobacillus casei | LACBCAS | Zygosaccharomyces species | ZYGSPE

## Table 16
14. Appendix E – BD Instrument Specifics
14.1. BACTEC MGIT 960
BACTEC MGIT 960 supports ASTM E_1394 Logical Protocol and ASTM E_1381 Physical Protocol.
14.1.1. Configurable Options
LIS Enable –acceptable values are ENABLE or DISABLE. Used to enable the LIS interface within the instruments. The instrument defaults to DISABLE.
Baud Rate – acceptable values are 1200, 2400, 4800, 9600, 19200. The instrument
defaults to 9600.
Data Bits – acceptable values are 7 and 8. The instrument defaults to 8.
Stop Bits – acceptable values are 1 and 2. The instrument defaults to 1.
Parity - acceptable values are ODD, EVEN, and NONE. The instrument defaults to NONE.
Upload Results – acceptable values are SOLICITED and UNSOLICITED. The instrument defaults to UNSOLICITED. In unsolicited mode the instrument automatically uploads results as tests are completed. The instrument in solicited mode waits for LIS to request the results.
Consumable Tracking – acceptable values include ENABLE or DISABLE. The instrument defaults to DISABLE. Enabling this option causes the instrument to automatically send a specimen’s test status to the LIS when a specimen has been scanned-in the instrument for the first time (ONGOING) or scanned-out of the instrument (REMOVED) (automatic transfer when UPLOAD_RESULT option set to UNSOLICITED).
Orphan Result – acceptable values include ENABLE and DISABLE. The instrument defaults to DISABLE. Enabling this option causes the instrument to report results for tests that do NOT have an accession number associated with them.
Packed Frames – acceptable values include ENABLE and DISABLE. The instrument defaults to ENABLE. Enabling this option allows message frames sent to the LIS to support multiple records per frame (“packed”). When Disabled, only one record is tranmitted per frame.
14.1.2. Message Content
Field List
The BACTEC MGIT 960 shall exchange fields with the LIS per the following table:
Header Record Field Name | ASTM Pos. | Direction
Sender Name | H, 5, 1, 1 | U
Version Number | H, 13, 1, 1 | U
Message Date/Time | H 14, 1, 1 | U
Order Record Field Name | ASTM Pos. | Direction
Accession Number | O, 3, 1, 1 | U
Isolate Number | O, 3, 1, 2 | U
Test ID | O, 5, 1, 4 | U

## Table 17
Result Record Field Name | ASTM Pos. | Direction
Result Type Code | R, 3, 1, 4 | U
Test/Consumable Sequence Number | R, 3, 1, 5 | U
Antibiotic | R, 3, 1, 6 | U
Antibiotic Concentration | R, 3, 1, 7 | U
Antibiotic Concentration Units | R, 3, 1, 8 | U
Test Status | R, 4, 1, 1 | U
Result Data 1
BACTEC MGIT 960 Growth Units | R, 4, 1, 2 | U
Result Data 3
AST susceptibility (Interpreted) | R, 4, 1, 4 | U
Preliminary/Final Status | R, 9, 1, 1 | U
Test Start Date/Time | R, 12, 1, 1 | U
Result/Status Date/Time | R, 13, 1, 1 | U
Instrument Type | R, 14, 1, 1 | U
Protocol Length | R, 14, 1, 3 | U
Instrument Number | R, 14, 1, 4 | U
Instrument Location | R, 14, 1, 5 | U
Request Record Field Name | ASTM Pos. | Direction
Request Starting Patient Id | Q, 3, 1, 1
Request Starting Accession Number | Q, 3, 1, 2 | D
Request Starting Sequence Number | Q, 3, 1, 3 | D
Request Ending Patient Id | Q, 4, 1, 1
Request Ending Accession Number | Q, 4, 1, 2
Request Ending Sequence Number | Q, 4, 1, 3
Request Test Id | Q, 5, 1, 1 | D
Request Test Status | Q, 5, 1, 2 | D
Request Instrument Type | Q, 5, 1, 3
Request Instrument Number | Q, 5, 1, 4
Request Result Qualifier | Q, 5, 1, 5 | D
Request Time Qualifier | Q, 6, 1, 1
Starting Date/Time | Q, 7, 1, 1 | D
Ending Date/Time | Q, 8, 1, 1 | D
Request Information Status Code | Q, 13, 1, 1
Terminator Record Field Name | ASTM Pos. | Direction
Termination Code | L, 3, 1, 1 | U/D
Field Descriptions
Header Fields
General
Patient Fields
BACTEC MGIT 960 does not support patient information and therefore always transmits an empty Patient record.
Order Fields
Accession Number (O, 3, 1, 1) – The unique alphanumeric string that identifies a specimen. This field can be up to 20 characters long. (e.g. Acc123).
Isolate Number (O, 3, 1, 2) – Specimen’s assigned isolate number.
Test Id (O, 5, 1, 4) –MGIT 960 transmits MGIT_960_GND for Growth and Detection specimens or MGIT_960_AST for AST sets.
Result Fields
Result Id Code (R, 3, 1, 4) – transmits GND_MGIT for growth and Detection specimens or AST_MGIT for AST sets.
Test Sequence Number (R, 3, 1, 5) – specimen’s assigned sequence number (barcode) that is always be 12 digits.
Drug Abbreviation (R, 3, 1, 6) – Antibiotic abbreviation used for a drug within an AST set.
Drug Concentration (R, 3, 1, 7) – Drug concentration used for a drug within an AST set.
Drug Units of Measurement (R, 3, 1, 8) – Drug units of measurement used for a drug within an AST set.
Test Status LIS Code (R, 4, 1, 1) – BACTEC MGIT 960 transmits the following specimen status:
INST_ONGOING INST_POSITIVE INST_NEGATIVE INST_COMPLETE
INST_REMOVED (for Ongoing Removed only)
INST_ERROR
It should be noted that Result records are not transmitted for G&D specimens that are in error. However, Result records are transmitted for AST set specimens that are in error because an AST set in error is considered a completed status. An AST set in error cannot return to an Ongoing state. A G&D specimen in error requires the users to resolve the before it returns to an Ongoing or Completed state.
Result Data (R, 4, 1, 2) – The BACTEC MGIT 960 instrument measures Mycobacterial growth. The instrument assigns a numerical value to quantify the amount of growth in a tube. This growth value is placed in the second component for any results generated by
this instrument, one of either GND_MGIT or AST_MGIT tests. This field is never more than 5 characters long.
AST SIR Status (R, 4, 1, 4) – The BACTEC MGIT 960 instrument transmits an ‘S’ for susceptible, ‘I’ for Indeterminate (not currently used), or an ‘R’ for resistant.
Result Status Code (R, 9, 1, 1) – Always a ‘P’ for Premliminary result.
Start Date/Time (R, 12, 1, 1) – This is the date and time that the test was first started or entered into an instrument. This field is formatted as described in the ASTM E_1394 specification in section 6.6.2.
The BACTEC MGIT 960 AST tests are run in reusable plastic carriers. The plastic carrier contains a sequence number that uniquely identifies the tube carrier. This is the sequence number transmitted with results up to the LIS. However, it is possible that a single plastic carrier could be used more than once for a particular access number. If the LIS intends to allow these carriers to be used more than once for the same access number, then the LIS should take note of the test start date/time. The test start date/time, with the sequence number, uniquely identifies a test and its result.
MGIT AST test information downloaded from the LIS is checked for this start date/time. If this field is filled, meaning that existing results are being copied back to the EpiCenter, the EpiCenter tries to use it to match an existing test in its database. New test orders should not fill this field. If the test sequence number and entry date/time don’t match an existing test, or are blank, then the EpiCenter considers it a new test order.
Test Result Date/Time (R, 13, 1, 1) – The instrument’s date and time that the specimen completed its test.
 For an INST_ONGOING specimen this field is empty.
 For an INST_POSITIVE specimen this is the time the specimen went positive.
 For an INST_NEGATIVE specimen this is the time the specimen went negative (calculated by adding the protocol length with the start time).
 For an INST_REMOVED specimen, this is the time the specimen was removed from the instrument.
 For an INST_COMPLETE specimen this is the time the AST set completed.
 For an INST_ERROR specimen this is the time the AST set went into error (this is considered a completed status).
Instrument Type (R, 14, 1, 1) – transmits a ‘MGIT960’.
Protocol Length (R, 14, 1, 3) – the specimen’s assigned protocol length.
Instrument Number (R, 14, 1, 4) –the user configured instrument number that ranges from 1 to 99.
Instrument Location (R, 14, 1, 5) – the specimen’s location within the instrument. For example, Drawer A, Row D, Column 12 would be represented as “A/D12”.
Request Fields
The ASTM 1394 protocol defines this as a Request Information Record message. A download Request record is also known as a “Query”. The only type of download “Query” request that the Firefly instrument accepts is a request for test results. A Terminator record with a termination code of “F” or “Q” is used to signal the last packet of a query response. If a query request was invalid, the instrument interface responds
with only the Terminator record with a termination code of “Q”. If the instrument interface can not locate any specimens in the active or history databases that meet the query criteria, it responds with only the Terminator record with a termination code of “F”. Note, that the BACTEC MGIT 960 instrument does NOT request data from an LIS.
A query contains a set of request parameters that are used to determine which specimen(s) to access. The following tables define the ASTM 1394 Request record fields that are transmitted from the LIS in an Download message and the MGIT 960 statuses that correspond to the StatusId field.
Accession Number (Q, 3, 1, 2) – The unique alphanumeric string that identifies a specimen. This field can be up to 20 characters long. (e.g. Acc123). The use of “ALL” or no definition indicates that the Accession # is not used as a search parameter in the query.
Sequence Number (Q, 3, 1, 3) – specimen’s assigned sequence number (barcode) that is always be 12 digits. The use of “ALL” or no definition indicates that the Sequence # is not used as a search parameter in the query.
Test ID (Q, 5, 1, 1) – The BACTEC 960 instrument accepts the test ids MGIT_960_GND, MGIT_960_AST, or ALL and return the test results for those test types currently in the instrument. The keyword ALL returns both GND and AST test results that are currently in the instrument.
Test Status Id (Q, 5, 1, 2) – The status of the specimens being requested. This field is only evaluated if the Test ID field contains valid data. No definition indicates that the specimen status is not used as a search parameter.
The following Test Statuses are used when requesting test results for specimens:
INST_ONGOING (G&D and AST) INST_POSITIVE INST_NEGATIVE INST_COMPLETE (G&D and AST)
Qualifier (Q, 5, 1, 5) – optional qualifier ‘MOD’ that indicates that the request is for specimen results that have been modified/changed since last request.
Beginning Time (Q, 7, 1, 1) – the beginning time and date used when requesting data that occurred within a specified time range.
End Time (Q, 8, 1, 1) – the ending time and date used when requesting data that occurred within a specified time range.
When a "Time Range" has been specified in a Query (Begin Time and/or End Time fields), the time field from the instrument's database that is used for comparison is dependent upon the Query's "TestId" field as follows,
INST_ONGOING - Time the specimen was scanned into instrument. INST_POSITIVE - Time the specimen went positive.
INST_NEGATIVE - Time the specimen went negative. INST_REMOVED - Time the specimen was removed. INST_COMPLETE - Time the AST set specimen completed. INST_ERROR - Time the AST set specimen went into error.
Terminator Fields
General
14.1.3. Example ASTM 1394 Protocol
The following are examples of various instrument message transmissions that conform to the Becton Dickinson Common LIS Interface using the ASTM 1394 protocol.
MGIT960 G&D specimen that has been entered into an instrument, but has NOT been tested:
H¦\^&¦¦¦Becton Dickinson¦7 Loveton Circle^Sparks^MD^21152¦¦410 316- 4000¦¦¦¦P¦V1.00¦19981019184200
P¦1
O¦1¦0970038018¦¦^^^MGIT_ 960_GND
R¦1¦^^^GND_MGIT^430109967349¦INST_ONGOING^0¦¦¦¦¦P¦¦¦¦¦MGIT960^^42^1^B/A01 L¦1¦N
MGIT960 G&D specimen that has been tested and is still active in the instrument:
H¦\^&¦¦¦Becton Dickinson¦7 Loveton Circle^Sparks^MD^21152¦¦ 410 316-4000¦¦¦¦P¦V1.00¦19981019184200
P¦1 O¦1¦0970038018¦¦^^^MGIT_960_GND
R¦1¦^^^GND_MGIT^430109967349¦INST_ONGOING^123¦¦¦¦¦P¦¦¦19960502115414¦ 19960507115414¦MGIT960^^42^1^B/A01
L¦1¦N
MGIT960 G&D specimen that has been tested POSITIVE:
H¦\^&¦¦¦Becton Dickinson¦7 Loveton Circle^Sparks^MD^21152¦¦ 410 316-4000¦¦¦¦P¦V1.00¦19981019184200
P¦1 O¦1¦0970038018¦¦^^^MGIT_960_GND
R¦1¦^^^GND_MGIT^430109967349¦INST_POSITIVE^1234¦¦¦¦¦P¦¦¦19960502115414¦ 19960509123456¦MGIT960^^42^1^B/A01
L¦1¦N
MGIT960 G&D specimen that has been tested NEGATIVE:
H¦\^&¦¦¦Becton Dickinson¦7 Loveton Circle^Sparks^MD^21152¦¦ 410 316-4000¦¦¦¦P¦V1.00¦19981019184200
P¦1 O¦1¦0970038018¦¦^^^MGIT_960_GND
R¦1¦^^^GND_MGIT^430109967349¦INST_NEGATIVE^12¦¦¦¦¦P¦¦¦19960502115414¦ 19960509123456¦MGIT960^^42^1^1B/A01
L¦1¦N
MGIT960 AST specimens that have been entered into an instrument, but has NOT been tested:
H¦\^&¦¦¦Becton Dickinson¦7 Loveton Circle^Sparks^MD^21152¦¦ 410 316-4000¦¦¦¦P¦V1.00¦19981019184200
P¦1 O¦1¦0970038018^1¦¦^^^MGIT_960_AST
R¦1¦^^^AST_MGIT^439309967349^EMB^5.0^ug/ml¦INST_ONGOING^0¦¦¦¦¦P¦¦¦¦¦ MGIT960^^^13^B/A02
R¦2¦^^^AST_MGIT^439309967349^RIF^1.0^ug/ml¦INST_ONGOING^0¦¦¦¦¦P¦¦¦¦¦ MGIT960^^^13^B/A03
L¦1¦N
MGIT960 AST specimens from an AST set that have been tested and are still active in the instrument:
H¦\^&¦¦¦Becton Dickinson¦7 Loveton Circle^Sparks^MD^21152¦¦ 410 316-4000¦¦¦¦P¦V1.00¦19981019184200
P¦1 O¦1¦0970038018^1¦¦^^^MGIT_960_AST
R¦1¦^^^AST_MGIT^439309967349^EMB^5.0^ug/ml¦INST_ONGOING^123¦¦¦¦¦P¦¦¦ 19960502115414¦19960502115414¦MGIT960^^^13^B/A02
R¦2¦^^^AST_MGIT^439309967349^RIF^1.0^ug/ml¦INST_ONGOING^12¦¦¦¦¦P¦¦¦ 19960502115414¦19960502115414¦MGIT960^^^13^B/A03
L¦1¦N
MGIT960 AST specimens that have completed testing:
H¦\^&¦¦¦Becton Dickinson¦7 Loveton Circle^Sparks^MD^21152¦¦ 410 316-4000¦¦¦¦P¦V1.00¦19981019184200
P¦1 O¦1¦0970038018^1¦¦^^^MGIT_960_AST
R¦1¦^^^AST_MGIT^439309967349^EMB^5.0^ug/ml¦INST_COMPLETE^1234^R¦¦¦¦¦P¦¦¦ 19960502115414¦19960502115414¦MGIT960^^^13^B/A02
R¦2¦^^^AST_MGIT^439309967349^RIF^1.0^ug/ml¦INST_COMPLETE^12^S¦¦¦¦¦P¦¦¦ 19960502115414¦19960502115414¦MGIT960^^^13^B/A03
L¦1¦N
A Request Record (Query) sent from an LIS to a MGIT960 instrument requesting all specimen results for a period of time.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19990208000001 Q¦1¦^ALL^ALL¦¦ALL¦¦19980502123000¦19980613123000
L¦1¦N
A Request Record (Query) sent from an LIS to a MGIT960 instrument requesting only ONGOING specimen results for a period of time.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19990208000001 Q¦1¦^ALL^ALL¦¦INST_ONGING¦¦19980502123000¦19980613123000
L¦1¦N
A Request Record (Query) sent from an LIS to a MGIT960 instrument requesting all specimen results which have been assigned the specified Accession #.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19990208000001 Q¦1¦^Access1234¦¦ALL
L¦1¦N
A Request Record (Query) sent from an LIS to a MGIT960 instrument requesting a specimen result which has been assigned the specified Sequence #.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19990208000001 Q¦1¦^^430109967349¦¦ALL
L¦1¦N
A Request Record (Query) sent from an LIS to a MGIT960 instrument requesting all specimen results which have been modified/changed since the last request.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19990208000001 Q¦1¦¦¦ALL^^^^MOD
L¦1¦N
A Request Record (Query) sent from an LIS to a MGIT960 instrument requesting all G&D specimen results that went Positive between January 2, 1999 at 15:34:00 and January 3, 1999 at 15:34:00.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦19990208000001 Q¦1¦¦¦MGIT_GND^INST_POSITIVE¦¦19990102123400¦19990103123400
L¦1¦N

## Table 18
14.2. BDProbeTec ET, Viper XTR, & Viper LT
BDProbeTec ET supports ASTM E_1394 Logical Protocol and ASTM E_1381 Physical Protocol.
BDProbeTec ET does not send results for specimens without accession or specimen numbers. As such the instrument does not allow LIS communication to be enabled unless the instrument Specimen Tracking/Accession Barcoding feature is enabled. The feature is enabled from a configuration screen.
14.2.1. Configurable Options
Comm Port Number – is a dedicated communication port on the instrument and therefore is not configurable.
Baud Rate – acceptable values are 2400, 4800, 9600, 14400, 19200, 38400. The instrument defaults to 9600. Viper LT will support Baud Rates of 300, 600, 1200, 1800, 2000, 2400, 3600, 4800, 7200, 9600, 14400, 19200, 28800, 38400, 57600,
115200
Data Bits – acceptable values are 7 and 8. The instrument defaults to 8. Stop Bits – acceptable values are 1 and 2. The instrument defaults to 1.
Parity - acceptable values are ODD, EVEN, and NONE. The instrument defaults to NONE.
Upload Mode – acceptable values are SOLICITED and UNSOLICITED. The instrument defaults to UNSOLICITED. The instrument in unsolicited mode uploads results as tests are completed. The instrument in solicited mode waits for the LIS to request the results.
Low Positive Status Upload – select 1 to enable the option or 0 to disable the option on BDProbeTec ET. Check to enable the option on Viper XTR. If the option is enabled, a MOTA value greater than or equal to 2,000 and less than 10,000 will produce an INST_LOW_POS Test Status LIS Code (R,4,1,1). Disabled is the default.
This feature is not supported for Viper LT.
Upload Algorithm Results – ( Viper XTR & Viper LT ) – Select check box to upload calculated algorithm value to the LIS.
Packed/Unpacked Frames – ( Viper XTR & Viper LT ) – Select checkbox to upload frames in Packed format.
14.2.2. Message Content
Field List
Header Record Field Name | ASTM Pos. | Direction
Sender Name | H, 5, 1, 1 | U
Version Number | H, 13, 1, 1 | U
Message Date/Time | H 14, 1, 1 | U

## Table 19
Order Record Field Name | ASTM Pos. | Direction
Accession Number | O, 3, 1, 1 | U
Test ID | O, 5, 1, 4 | U
Specimen Action Code | O, 12, 1, 1 | U
Result Record Field Name | ASTM Pos. | Direction
Result Type Code | R, 3, 1, 4 | U
Test/Consumable Sequence Number | R, 3, 1, 5 | U
Test Status | R, 4, 1, 1 | U
Result Data 1
ProbeTecET/Viper XTR/Viper LT Alg.Results | R, 4, 1, 2 | U
Test Start Date/Time | R, 12, 1, 1 | U
Result/Status Date/Time | R, 13, 1, 1 | U
Instrument Type | R, 14, 1, 1 | U
Media/Assay Type | R, 14, 1, 2 | U
Instrument Number | R, 14, 1, 4 | U
Instrument Location | R, 14, 1, 5 | U
QC Type | R, 14, 1, 6 | U
QC Kit Lot Number | R, 14, 1, 7 | U
Request Record Field Name | ASTM Pos. | Direction
Request Starting Patient Id | Q, 3, 1, 1
Request Starting Accession Number | Q, 3, 1, 2
Request Starting Sequence Number | Q, 3, 1, 3
Request Ending Patient Id | Q, 4, 1, 1
Request Ending Accession Number | Q, 4, 1, 2
Request Ending Sequence Number | Q, 4, 1, 3
Request Test Id | Q, 5, 1, 1 | D
Request Test Status | Q, 5, 1, 2
Request Instrument Type | Q, 5, 1, 3
Request Instrument Number | Q, 5, 1, 4
Request Result Qualifier | Q, 5, 1, 5 | D
Request Time Qualifier | Q, 6, 1, 1
Starting Date/Time | Q, 7, 1, 1
Ending Date/Time | Q, 8, 1, 1
Request Information Status Code | Q, 13, 1, 1
Terminator Record Field Name | ASTM Pos. | Direction
Termination Code | L, 3, 1, 1 | U/D
Field Descriptions
Header Fields
General
Patient Fields
BDProbeTec ET does not support patient information and therefore always transmits an empty Patient record.
P¦1
Order Fields
Accession Number (O, 3, 1, 1) – The unique alphanumeric string that identifies a specimen. This field can be up to 20 characters long. (e.g. Acc123) This is a required field for processing specimen and test information. The instrument excludes the following special characters from being entered in the accession / specimen number field and therefore never transmits them: * ? [ ] #. Viper LT excludes the following additional characters: !¦.
Test Id (O, 5, 1, 4) –BDProbeTec ET transmits BD_PROBETEC_GND for the test id code field.
Specimen Action Code (O, 12, 1, 1) - In the special case when no accession / specimen number is transmitted then the Specimen Action Code is designated ‘Q’ as an indicator to treat the result as a QC result.
P¦1
O¦1¦Accession_789 ¦ ¦ ^ ^ ^BD_PROBETEC_GND
P¦1
O¦1¦ ¦ ¦ ^ ^ ^BD_PROBETEC_GND¦ ¦ ¦ ¦ ¦ ¦ ¦Q
Result Fields
Result Id Code (R, 3, 1, 4) – BDProbeTec ET always transmits ‘GND_PROBETEC’ for the result id code.
Test Sequence Number (R, 3, 1, 5) – Designates the assay kit lot number and is always 7 digits.
Test Status LIS Code (R, 4, 1, 1) – BDProbeTec ET transmits the following specimen status:
INST_POSITIVE
INST_LOW_POS (If enabled, see Configurable Options. Not sent by Viper LT)
INST_NEGATIVE INST_INDETERMINATE
INST_EQUIVOCAL (Not sent for Viper LT)
The instrument transmits the following QC status:
INST_QC_PASS INST_QC_FAIL
Algorithmic Results (R, 4, 1, 2) – If enabled, see Configurable Options.
Test Start Date/Time (R, 12, 1, 1) – BDProbeTec ET transmits the instrument’s date and time that the specimen started its test. Time zone is not included.
Test Result Date/Time (R, 13, 1, 1) – BDProbeTec ET transmits the instrument’s date and time that the specimen ended its test. Time zone is not included. Typically, the test result date and time are 1 hour after the test start date and time.
Instrument Type (R, 14, 1, 1) – BDProbeTec ET transmits a ‘PROBETEC’ for its instrument type.
Media (Assay) Type (R, 14, 1, 2) – BDProbeTec ET transmits the following assay types:
1 – CT assay (Viper XTR)
2 – GC assay (Viper XTR)
3 – Intentionally skipped
4 – DTB assay
5 – ctb assay 6 – mac assay 7 – kan assay
8 – Intentionally skipped
9 – CF assay 10 – LP assay 11 – MP assay 12 – tCF assay 13 – tMP assay
14 – dCT assay (Viper XTR, Viper LT)
15 – dGC assay (Viper XTR, Viper LT)
16 – SCV assay
17 – HSV1 assay (Viper XTR)
18 – HSV2 assay (Viper XTR)
19 – HPV assay (Viper XTR – HR for Viper LT)
20 – TV assay (Viper XTR)
200 – HPV 16 (Viper LT)
201 – HPV 18 (Viper LT)
202 – HPV 45 (Viper LT)
203 – HPV 31 (Viper LT)
204 – HPV 51 (Viper LT)
205 – HPV 52 (Viper LT)
206 – HPV 33/58 (Viper LT)
207 – HPV 56/59/66 (Viper LT)
208 – HPV 35/39/68 (Viper LT)
209 – GT (Viper LT)
The list of assay types will expand as further assays are supported on the BDProbeTec ET instrument.
Protocol Length (R, 14, 1, 3) – BDProbeTec ET currently has a 1 hour test time and therefore does not transmit a protocol length.
Instrument Number (R, 14, 1, 4) – BDProbeTec ET transmits the user configured instrument number that ranges from 1 to 99.
Instrument Location (R, 14, 1, 5) – BDProbeTec ET transmits a plate number from 1 to 30, followed by a 96 microwell location of row (A-H), then column (1-12)
Viper XTR transmits a plate number from 1 to 60, followed by a 96 microwell location of row (A-H), then column (1-12).
Viper LT transmits a rack number from 1 to 30, followed by a 96 microwell location of row (A-H), then column (1-12).
QC Type (R, 14, 1, 6) – BDProbeTec ET transmits a ‘+’ or ‘-‘ if the parent Order record designates a QC result to indicate what type of QC, positive or negative QC.
QC Kit Lot Number (R, 14, 1, 7) – BDProbeTec ET transmits a value only if parent order record is designated as a QC test.
P¦1
O¦1¦Accession_789 ¦ ¦ ^ ^ ^BD_PROBETEC_GND
R¦1¦ ^ ^ ^ GND_PROBETEC ^000120299¦INST_NEGATIVE¦ ¦ ¦ ¦ ¦ ¦ ¦
¦19981019153400¦19981019163400¦ PROBETEC ^1^^3^14A10
P¦1
O¦1¦ ¦ ¦ ^ ^ ^BD_PROBETEC_GND¦ ¦ ¦ ¦ ¦ ¦ ¦Q
R¦1¦ ^ ^ ^ GND_PROBETEC ^000120299¦INST_QC_PASS¦ ¦ ¦ ¦ ¦ ¦ ¦
¦19981019153400¦19981019163400¦ PROBETEC ^1^^3^14H12^+^000031500
Request Fields
The ASTM 1394 protocol defines this as a Request Information Record message. A download Request record is also known as a “Query”. The only type of download “Query” request that the BD ProbeTec ET instrument accepts is a request for test results. The BD Probetec ET instrument will accept "query" requests in either the SOLICITED or UNSOLICITED settings for the Upload Mode as described above under Configurable Options. Note, that the BD PROBETEC ET instrument does NOT request data from an LIS.
For an ALL request, Viper LT will send up to 30 of the most recently completed racks.
Test ID (Q, 5, 1, 1) – BDProbeTec ET accepts the following test name: BD_PROBETEC_GND. The keyword ALL can also be used to return all of the test results currently in the instrument.
Q¦1¦ ¦ ¦BD_PROBETEC_GND
Q¦1¦ ¦ ¦ALL
Result Qualifier (Q, 5, 1, 5) – BDProbeTec ET accepts the following result qualifiers:
MOD – All available results that have not been previously transmitted to the LIS. Q¦1¦ ¦ ¦BD_PROBETEC_GND^^^^MOD
Terminator Field
General
14.2.3. Example ASTM 1394 Protocol
The following are examples of valid requests for test results from the LIS to the BD PROBETEC ET instrument that conform to the Becton Dickinson Common LIS Interface using the ASTM 1394 protocol:
A Request Record (Query) sent from an LIS to a BD PROBETEC ET instrument r equesting all test results currently in the instrument.
For an ALL request, Viper LT will send up to 30 of the most recently completed racks.
H¦\^&
Q¦1¦¦¦ALL

## Table 20
L¦1¦N
A Request Record (Query) sent from an LIS to a BD PROBETEC ET instrument requesting all available test results that have not been previously transmitted to the LIS.
H¦\^&
Q¦1¦¦¦ BD_PROBETEC_GND^^^^MOD
L¦1¦N
14.2.4. Unsolicted Demographics Download (Viper LT)
Viper LT shall accept patient demographic fields from the LIS per the following tables:
Header Record Field Name | ASTM Pos. | Direction
Sender Name | H, 5, 1, 1 | U/D
Version Number | H, 13, 1, 1 | U/D
Message Date/Time | H 14, 1, 1 | U/D
Patient Record Field Name | ASTM Pos. | Direction
Patient ID | P, 4, 1, 1 | D
Date of Birth | P, 8, 1, 1 | D
Patient Sex | P, 9, 1, 1 | D
Hospital Service | P, 33, 1, 1 | D
Hospital Client | P, 34, 1, 1 | D
Order Record Field Name | ASTM Pos. | Direction
Accession Number | O, 3, 1, 1 | D
Collection Date Time | O, 8, 1, 1 | D
Terminator Record Field Name | ASTM Pos. | Direction
Termination Code | L, 3, 1, 1 | U/D
14.2.4.1.Patient Fields
Patient ID (P, 4, 1, 1) - Identifier that uniquely identifies a patient. This is a required field for patient demographic downloads. This field can be up to 16 characters long (e.g. PatId123). If the field is blank, the record is ignored.
Date of Birth (P, 8, 1, 1) – patient date of birth formatted as described in ASTM E_1394 section 6.6.2. This field is optional.
Patient Sex (P, 9, 1, 1) – There are 3 predefined values accepted for this (M)ale, (F)emale, and (U)nspecified. If the value does not match the code for either Male or Female then the sex is set to Unspecified. This field is optional.
Hospital Service (P, 33, 1, 1) – An acronym or abbreviation for the ward or hospital department involved in caring for this patient (e.g. ER may represent the Emergency Room). This field can be up to 20 characters long.
Hospital Client (P, 34, 1, 1) – This is the facility that the specimen came from, for instance Smith Radiology or Central Labs. This field is optional. This field can be up to 20 characters long.
14.2.4.2.Order Fields
Accession Number (O, 3, 1, 1) - Maximum field length: 20 characters. A unique alphanumeric string that identifies a specimen. If the field is blank, the record is ignored.
Collection Date Time (O, 8, 1, 1) Date and time that the specimen was collected from the patient, formatted as described in ASTM E_1394 section 6.6.2.
H¦\^&¦¦¦¦¦¦¦¦¦¦¦¦20130215184200 P¦1¦¦PatId123¦¦^^^^¦¦19651029¦M¦¦^^^^¦¦¦¦¦¦
¦¦¦¦¦¦¦¦¦¦¦¦¦¦ER¦St. Josephs Hospital O¦1¦Acc123¦¦^^^^¦¦¦20130214023300¦¦¦¦¦¦¦ ¦^¦¦^^¦¦¦¦¦¦¦¦¦¦ L¦1¦N
14.3. BD Phoenix 100
BD Phoenix 100 supports ASTM E_1394 Logical Protocol and ASTM E_1381 Physical Protocol.
14.3.1. Configurable Options
The BD Phoenix 100 provides 3 groups of Configurable Options for customizing the LIS Interface.
General Options Results Upload Options Communications Setup
The instrument also provides an Organism Configuration Screen and an Antimicrobial Configuration Screen for editing LIS Codes.
14.3.1.1.General Options
Send Interpretation Results
Acceptable values are CHECKED and UNCHECKED. When CHECKED, final SIR values are included in the AST result record. The instrument defaults to CHECKED.
Unsolicited Queries
Acceptable values are CHECKED and UNCHECKED. When CHECKED, the instrument shall issue a query to the LIS at the time the panel is physically placed in the instrument if the panel is missing an organism id. The instrument defaults to UNCHECKED.
Send When Placed In Instrument
Acceptable values are CHECKED and UNCHECKED. When CHECKED, the instrument shall send a result upload to the LIS at the time the panel is physically placed in the instrument. The instrument defaults to CHECKED.
14.3.1.2.Results Upload Options
There are five results upload options, only one of which can be CHECKED at a time.
Solicited
Acceptable values are CHECKED and UNCHECKED. When CHECKED the instrument waits for the LIS to request the results The instrument defaults to UNCHECKED.
Send On Finalization
Acceptable values are CHECKED and UNCHECKED. When CHECKED the instrument shall send results to the LIS when the panel is finalized or a modification has been made to a finalized panel. The instrument defaults to UNCHECKED.
Send On Completion
Acceptable values are CHECKED and UNCHECKED. When CHECKED the instrument shall send results to the LIS when the panel’s status transitions to COMPLETE or a modification has been made to a COMPLETE panel . The instrument defaults to CHECKED.
Send as Available
Acceptable values are CHECKED and UNCHECKED. When CHECKED the instrument shall send results to the LIS when the panel’s status transitions to COMPLETE or the panel has partial results. Partial results are when at least one MIC value is determined for an AST or ID/AST panel or the organism ID is determined for an ID/AST panel. The instrument defaults to UNCHECKED.
Send at Fixed Time
Acceptable values are CHECKED and UNCHECKED. When CHECKED the instrument shall send results to the LIS when the panel’s status transitions to COMPLETE and send panels with partial results to the LIS at the Fixed Time. When the Fixed Time is reached, any panel that is ONGOING and has partial results shall be uploaded to the LIS with a Test Status of INST_PARTIAL_COMPLETE. The instrument defaults to UNCHECKED.
14.3.1.3.Communications Setup
NOTE: Modifying any of the following options will result in the instrument rebooting.
Baud Rate
Acceptable values are 2400, 4800, 9600, 14400,19200,38400. The instrument
defaults to 9600.
Data Bits
Acceptable values are 7 and 8. The instrument defaults to 8.
Parity
Acceptable values are ODD, EVEN, and NONE. The instrument defaults to NONE.
Stop Bits
Acceptable values are 1 and 2. The instrument defaults to 1.
Packed Frames
Acceptable values are YES and NO. Selecting YES allows message frames sent to the LIS to support multiple records per frame (“packed”). When NO is selected, only one record is transmitted per frame. The instrument defaults to YES.
14.3.1.4.Organism and Antimicrobial Configuration Screens
Except for antimicrobials listed in special messages and BDXpert Rules within comment fields, all referrals to organism IDs and antimicrobials in LIS communications use the LIS Codes stored in the instrument which are set to default values per Appendix C and Appendix D. These codes can be edited to a different value.
14.3.2. Modes of Communications
Download refers to communications from the LIS to the instrument and upload refers to communications from the instrument to the LIS. NOTE: The total size of a message downloaded to the instrument should not exceed 100k characters.
Note: New in version V5.02, results for orphan panels and QC panels are not uploaded to the LIS unless the panels satisfy the criteria of an LIS query.
The BD Phoenix 100 shall support the following modes of communications with an LIS:
Result Uploads
The BD Phoenix 100 shall upload results according to the CHECKED Results Upload Option. For AST sets, several result records shall be sent in a single message. Each result record shall report the results for a single drug in the AST set. For ID/AST combo tests, an ID result record is sent as the first result record, and the AST result records follow
Query Uploads
The BD Phoenix 100 shall upload queries if the Unsolicited Queries Option is CHECKED.
Order Downloads
The BD Phoenix 100 instrument shall accept order downloads. All orders must include either a Test ID or a Sequence Number, a valid Accession number and a valid Isolate Number. An order consisting of a valid Consumable Sequence Number and valid Accession\Isolate Number shall be saved to the instrument database as a PENDING panel. Orders without the Consumable Sequence Number shall be saved in the instrument, BUT must be associated to a Consumable Sequence Number via the LIS Login Screen prior to being saved to the instrument database. Orders with invalid data shall be rejected by the instrument. See the Phoenix User Manual (RE: LIS Order Cancelled Messages) for more information regarding cancelled LIS orders.
If an order is received with an organism ID that is either unknown to the instrument or is an unclaimed organism, the order will still be saved to the database or saved for future association with a Consumable Sequence Number, BUT the organism ID will be replaced with Invalid Organism Received.
If an order is downloaded with a sequence number that already exists in the instrument and all populated order fields are valid, the panel record is populated/overwritten with the order fields. Modifications shall not be accepted for Finalized panels.
An order field with all spaces will be treated as “missing”. Leading or trailing spaces are trimmed.
Query Downloads
The BD Phoenix 100 instrument shall accept download queries. The instrument shall accept queries on the following fields: Result Qualifier, Start Time, End Time, Request Time Qualifier, Sequence Number, Test ID, Status and Accession. Queries with invalid data shall be rejected by the instrument. See the Phoenix User Manual (RE: LIS Query Assembly Messages) for more information regarding invalid queries.

## Table 21
14.3.3. Message Content
The BD Phoenix 100 shall exchange fields with the LIS per the following table:
Field List
The BD Phoenix 100 shall exchange fields with the LIS per the following table:
Header Record Field Name | ASTM Pos. | Direction
Sender Name | H, 5, 1, 1 | U
Version Number | H, 13, 1, 1 | U
Message Date/Time | H 14, 1, 1 | U
Order Record Field Name | ASTM Pos. | Direction
Accession Number | O, 3, 1, 1 | U/D
Isolate Number | O, 3, 1, 2 | U/D
Organism | O, 3, 1, 3 | U/D
Test ID | O, 5, 1, 4 | U/D
Test/Consumable Sequence Number | O, 5, 1, 5 | U/D
Priority | O, 6, 1, 1 | U/D
Report Type | O, 26, 1, 1 | U
Result Record Field Name | ASTM Pos. | Direction
Result Type Code | R, 3, 1, 4 | U
Test/Consumable Sequence Number | R, 3, 1, 5 | U
Antibiotic | R, 3, 1, 6 | U
Test Status | R, 4, 1, 1 | U
Result Data 1
AST MIC for AST MIC Result | R, 4, 1, 2 | U
Organism ID for ID Result | R, 4, 1, 2 | U
Result Data 2
AST susceptibility (Final) | R, 4, 1, 3 | U
Profile Number for ID Test | R, 4, 1, 3 | U
Result Data 3
Resistance Marker 1 | R, 4, 1, 4 | U
Result Data 4
Resistance Marker 2 | R, 4, 1, 5 | U
Result Data 5
Resistance Marker 3 | R, 4, 1, 6 | U
Result Data 6
Resistance Marker 4 | R, 4, 1, 7 | U
Result Data 7
Resistance Marker 5 | R, 4, 1, 8 | U
Resistance Marker 6 | R, 4, 1, 9 | U
Resistance Marker 7 | R, 4, 1, 10 | U
Resistance Marker 8 | R, 4, 1, 11 | U
Resistance Marker 9 | R, 4, 1, 12 | U
Resistance Marker 10 | R, 4, 1, 13 | U
Preliminary/Final Status | R, 9, 1, 1 | U
Test Start Date/Time | R, 12, 1, 1 | U
Result/Status Date/Time | R, 13, 1, 1 | U
Test Complete Date/Time | R, 13, 2, 1 | U
Instrument Type | R, 14, 1, 1 | U
Instrument Number | R, 14, 1, 4 | U
Instrument Location | R, 14, 1, 5 | U

## Table 22
Comment Record Field Name | ASTM Pos. | Direction
Comment Text | C, 4, 1, 1 | U
Comment Type | C, 5, 1, 1 | U
Request Record Field Name | ASTM Pos. | Direction
Request Starting Accession Number | Q, 3, 1, 2 | U/D
Request Starting Sequence Number | Q, 3, 1, 3 | U/D
Request Test Id | Q, 5, 1, 1 | D
Request Test Status | Q, 5, 1, 2 | D
Request Result Qualifier | Q, 5, 1, 5 | D
Request Time Qualifier | Q, 6, 1, 1 | D
Starting Date/Time | Q, 7, 1, 1 | D
Ending Date/Time | Q, 8, 1, 1 | D
Terminator Record Field Name | ASTM Pos. | Direction
Termination Code | L, 3, 1, 1 | U/D
Field DescriptionsPatient Fields
BD Phoenix 100 does not support patient demographics and therefore shall send an empty Patient record.
Order Fields
Accession Number (O, 3, 1, 1) - Maximum field length: 20 characters. A unique alphanumeric string that identifies a specimen.
Isolate Number (O, 3, 1, 2) - Maximum field length: 2 characters. Allowable range: 1-
20. The number associated with an isolate.
Organism (O, 3, 1, 3) - Maximum field length: 20 characters. LIS code of the organism ID. The organism ID for an ID panel or an ID/AST panel is the User ID if it exists or if a User ID does not exist and the panel has a single Instrument ID it is the Instrument Organism ID. The organism ID for an AST panel is the User ID if it exists or if a User ID does not exist and the panel has been auto associated it is the Instrument Organism ID. Organism IDs and their corresponding LIS codes can be viewed/edited in the instrument’s Organism Configuration Screen
For QC Panels the organism ID shall be the Test Strain.
Test ID (O, 5, 1, 4) - Maximum field length: 20 characters. Populated with the panel format abbreviation as shown on the panel’s barcode label. Appendix B contains a sample list.
Test/Consumable Sequence Number (O, 5, 1, 5) - 12 digit numeric string identifying the instrument consumable.
Priority (O, 6, 1, 1) - Maximum field length: 1 character. This field is defined in the ASTM E_1394 specification. Currently only the values ‘A’ for Critical, ‘R’ for Normal and blank for default normal or no change are supported.
Report Type (O, 26, 1, 1) – This field is defined in the ASTM E_1394 specification. When a cancelled order is sent back to the LIS the report type is ‘X’.
Result Fields
The BD Phoenix 100 uploads shall consist of AST results and ID results. The information contained in the generic “Data Fields” depends on the type of result being reported. AST Result records contain antimicrobial and susceptibility information. ID Result records contain organism and Resistance Marker information.
Result ID Code (R, 3, 1, 4) A BD defined code that indicates the type of information in the result record. The keyword AST_MIC indicates results for a MIC based Antimicrobial Susceptibility Test and the keyword ID indicates identification results.
Test/Consumable Sequence Number (R, 3, 1, 5) – 12 digit numeric string identifying the instrument consumable. Antimicrobial (R, 3, 1, 6) (AST Results) Populated with the LIS code corresponding to the antimicrobial. Antimicrobials and their corresponding LIS codes can be viewed/edited in the instrument’s Antimicrobial Configuration Screen. This field can be up to 20 characters long.
Test Status LIS Code (R, 4, 1, 1) For non QC panels this field is populated with one of the statuses listed in Appendix B. For QC panels that are ONGOING or PENDING , the status is reported as for non QC panels, however, for COMPLETED panels the status is INST_QC_PASS or INST_QC_FAIL.
Result Data Field 1
The contents of the Result Data fields vary depending on the type of result being reported, as described below.
Minimum Inhibitory Concentration (R, 4, 1, 2) (MIC based AST Result) For AST tests that produce MIC values; this field contains the minimum antibiotic concentration that inhibits growth of the tested organism. This field can be up to 20 characters long and is assumed to contain a number in ug/ml.
This field may contain signed MIC values such as “<=4” or “>8”, MIC values for compound antimicrobials such as “0.5/4”or “<=0.5/16” as well as the following:
“?” - Ongoing
“C” - Antimicrobial is Rapid Completed (For nonQC tests in V4.01 and later)
“X” - Error
Note: For QC panels, MIC values of X will be uploaded to the LIS for V4.01. The uploaded MIC values for QC panels are ‘blank’ in previous software versions.
Organism (R, 4, 1, 2) (ID Results Only) LIS code of the organism ID. The organism ID for an ID panel or an ID/AST panel is the User ID if it exists or if a User ID does not exist and the panel has a single Instrument ID it is the Instrument Organism ID. The organism ID for an AST panel containing a resistance marker, otherwise the organism ID is only included in the order record. Organism IDs and their corresponding LIS codes can be viewed/edited in the instrument’s Organism Configuration Screen. This field can be up to 20 characters long.
Result Data Field 2
AST Susceptibility, Final (R, 4, 1, 3) ( AST results only) This is a code that indicates the organism’s susceptibility to the antibiotic specified in this record. The acceptable values for this field are, blank, S, I, R, N, X representing MIC is pending, susceptible, intermediate, resistant, not susceptible and error respectively.
Result Data Field 3
Resistance Marker 1 (R, 4, 1, 4) (ID results) This represents a Resistance Marker associated with the organism identified in the ID result record. This field can be up to 8 characters long. Examples of Resistance Markers and their corresponding LIS codes are listed below.
Extended Spectrum B-Lactamase RM_ESBL Methicillin Resistant Staphylococcus RM_MRSA Streptomycin HLAR RM_HLSR B-Lactamase in Gram Positive Cocci RM_GP_BL Vancomycin Resistant Enterococci RM_VRE
Gentamicin HLAR RM_HLGR
High Level Kanamycin Resistant RM_HLKR
Streptococcus MLSb Phenotype RM_MLSB Streptococcus Macrolide Efflux Phenotype RM_MEFF High Level Penicillin Resistant S. Pneumoniae RM_HLPRSP Low Level Penicillin Resistant S. Pneumoniae RM_LLPRSP High Level Mupiricin Resistant Staphylococcus RM_HLMUP
Result Data Field 4
Resistance Marker 2 (R, 4, 1, 5) (ID results) This represents a Resistance Marker associated with the organism identified in the ID result record. This field can be up to 8 characters long. Examples of Resistance Markers and their corresponding LIS codes are listed under Result Data Field 3.
Result Data Field 5
Resistance Marker 3 (R, 4, 1, 6) (ID results) This represents a Resistance Marker associated with the organism identified in the ID result record. This field can be up to 8 characters long. . Examples of Resistance Markers and their corresponding LIS codes are listed under Result Data Field 3.
Resistance Marker 4-10 (R, 4, 1, 7-13) (ID results) This represents a Resistance Marker associated with the organism identified in the ID result record. This field can be up to 8 characters long. . Examples of Resistance Markers and their corresponding LIS codes are listed under Result Data Field 3.
Preliminary/Final Status (R, 9, 1, 1) This field contains either a “P” for preliminary status or an “F” for final status, which occurs when the panel is finalized at the instrument.
Start Date/Time (R, 12, 1, 1) This is the date and time that the panel was physically placed in the instrument. This field is represented in the YYYYMMDDHHMMSS format.
Result/Status Date/Time (R, 13, 1, 1) This is the date and time that the panel’s status transitioned to COMPLETE. This field is represented in the YYYYMMDDHHMMSS format.
Test Complete Date/Time (R, 13, 2, 1) (ID results) This is the date and time that the ID panel obtained an ID. (AST results) This is the date & time that a drug obtained a valid MIC. This field is represented in the YYYYMMDDHHMMSS format.
Instrument Type (R, 14, 1, 1) Identifies the instrument with the keyword PHOENIX.
Instrument Number (R, 14, 1, 4) Instrument number between 1 and 99.
Instrument Location (R, 14, 1, 5) Indicates the panel position inside the instrument as a 3 character string represented in TierStation format. Tier values are A-D and station values are 1-25. This field is empty for PENDING panels.
Request FieldsStarting Access Number (Q, 3, 1, 2) - Maximum field length: 20 characters. This field should be populated if the request is for either a single specimen or for a range of specimens. If accession number is not to be used as a search criterion for requested data, then this field can be left blank. Populating this field with the keyword ALL queries for all panels in the instrument database that meet the remaining criteria.
Starting Sequence Number (Q, 3, 1, 3) - Required field length: 12 digit consumable ID value. This field should be populated if the request is for a single panel. If sequence number is not to be used as a search criterion for requested data, then this field can be left blank. Populating this field with the keyword ALL queries for all panels in the instrument database that meet the remaining criteria.
Test ID (Q, 5, 1, 1) - Maximum field length: 20 characters. This field should be populated with the test type (ID, AST_MIC). The BD Phoenix 100 accepts the keywords ID and AST_MIC as valid test types as well as the pre-defined test names. See appendix B for examples. Populating this field with the keyword ALL queries for all panels in the instrument database that meet the remaining criteria.
Test Status (Q, 5, 1, 2) The BD Phoenix 100 instrument accepts queries for panels with a status as follows:
 INST_ONGOING, panels with an onging status that may or may not yet contain results.
 INST_COMPLETE, panels with a complete, rapid complete, in attention ignored, in attention complete, QC pass, or QC fail status
 INST_RAPID_COMPLETE, panels with a rapid complete status
 INST_PARTIAL_COMPLETE, panels with an onging status that have obtained results.
 INST_QC_PASS, QC panels with a passed QC status
 INST_QC_FAIL, QC panels with a failed QC status
Result Qualifier (Q, 5, 1, 5) The only acceptable value for this field is the keyword
MOD to request only results that have been modified since the last request.
Time Qualifier (Q, 6, 1, 1) - Maximum field lenth: 1 character. The acceptable values for this field are ‘S’ which indicate the following date time fields refer to test start times or ‘R’ which indicate the following date time fields can refer to result/status times.
Starting Date/Time (Q, 7, 1, 1) The starting date/time (inclusive) of the data requested represented in the YYYYMMDDHHMMSS format.
Ending Date/Time (Q, 8, 1, 1) The ending date/time (inclusive) of the data requested represented in the YYYYMMDDHHMMSS format.
Comment Fields
The BD Phoenix 100 sends comment strings for BdXpert rule text and special messages.
The following is an example of a comment record containing a rule where <310> is the rule number and (AMX, ATM, CAZ) is a list of the affected antimicrobials.
C¦1¦ ¦<310> BDXpert Rule 310 text.(AMX, ATM, CAZ)¦E
Special Messages are messages that indicate a special condition for a BD Phoenix 100 panel.
The following is an example of a comment record containing a special message with a list of the affected antimicrobials.
C¦1¦ ¦ Special message text.(AM, CAZ)¦T
Comment Text (C, 4, 1, 1) – The text to be used for the patient, specimen or isolate comment. This field can be up to 1600 characters long.
Comment Type (C, 5, 1, 1) – The type of comment record. This field should be an E, or T for BDXpert rule, or special message comments respectively.
Terminator Field
General
14.3.4. Example ASTM 1394 Protocol
Result Upload Examples
Example of a rejected order
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20031110093857 P¦1
O¦1¦~!@#$^1¦¦^^^CT04N¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦X
C¦1¦¦LIS Order Cancelled: Invalid Accession Number Field¦T L¦1¦N
Example of a “Send When Placed in Instrument” result upload
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20031110101102 P¦1
O¦1¦IDAST 1^1¦¦^^^GNIDAST 951^429510000001¦R R¦1¦^^^ID^429510000001¦INST_ONGOING¦¦¦¦¦P¦¦¦20031110101102¦¦PHOENIX^^^1^C7
R¦2¦^^^AST_MIC^429510000001^AM¦INST_ONGOING^?^ ¦¦¦¦¦P¦¦¦ 20031110101102¦¦PHOENIX^^^1^C7
R¦3¦^^^AST_MIC^429510000001^AN¦INST_ONGOING^?^ ¦¦¦¦¦P¦¦¦ 20031110101102¦¦PHOENIX^^^1^C7
R¦4¦^^^AST_MIC^429510000001^AMC¦INST_ONGOING^?^ ¦¦¦¦¦P¦¦¦ 20031110101102¦¦PHOENIX^^^1^C7
L¦1¦N
Example of a result upload for an ID panel
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20031110102747 P¦1
O¦1¦ID 1^1^ALCPIE¦¦^^^GNID 952^429520000001¦R
R¦1¦^^^ID^429520000001¦INST_COMPLETE^ALCPIE¦¦¦¦¦P¦¦¦20031110101102¦ 20031110102747\20031110102747¦PHOENIX^^^1^D7
L¦1¦N
Example of result uploaded to an LIS with the Send Interpretation Results option UNCHECKED
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20031110114736 P¦1
O¦1¦IDAST 1^1^SHISPE¦¦^^^GNIDAST 951^429510000001¦R
R¦1¦^^^ID^429510000001¦INST_COMPLETE^SHISPE¦¦¦¦¦P¦¦¦20031110112648¦ 20031110114736\20031110114735¦PHOENIX^^^1^A7
R¦2¦^^^AST_MIC^429510000001^AM¦INST_COMPLETE^<=4^S¦¦¦¦¦P¦¦¦ 20031110112648¦20031110114736\20031110114735¦PHOENIX^^^1^A7
R¦3¦^^^AST_MIC^429510000001^AN¦INST_COMPLETE^<=4^S¦¦¦¦¦P¦¦¦ 20031110112648¦20031110114736\20031110114735¦PHOENIX^^^1^A7
R¦4¦^^^AST_MIC^429510000001^AMC¦INST_COMPLETE^<=4/2^S¦¦¦¦¦P¦¦¦ 20031110112648¦20031110114736\20031110114735¦PHOENIX^^^1^A7
L¦1¦N
Example of a result upload for an ID/AST panel with special messages, resistance markers and the Send Interpretation Results option UNCHECKED
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20031110150803
P¦1
O¦1¦ABC^1^ENTCFAA¦¦^^^CT01P^424940000029¦R
C¦1¦¦The MICs for this antibiotic and species combination are not reported by the Phoenix system. An alternate method should be utilized.(RA)¦T
C¦2¦¦Fill failures were detected with this antibiotic and no MIC can be determined. The isolate should be retested.(CF)¦T
R¦1¦^^^ID^424940000029¦INST_IN_ATTN_COMPLETE^ENTCFAA^^RM_VRE^RM
_HLSR^RM_HLGR¦¦¦¦¦P¦¦¦20031110145645¦20031110150803¦PHOENIX^^^1^C5
R¦2¦^^^AST_MIC^424940000029^AM¦INST_IN_ATTN_COMPLETE^>32¦¦¦¦¦P¦¦¦200311 10145645¦20031110150803\20031110150802¦PHOENIX^^^1^C5
R¦3¦^^^AST_MIC^424940000029^AZM¦INST_IN_ATTN_COMPLETE^>8¦¦¦¦¦P¦¦¦20031 110145645¦20031110150803\20031110150802¦PHOENIX^^^1^C5
R¦4¦^^^AST_MIC^424940000029^CAZ¦INST_IN_ATTN_COMPLETE^>64¦¦¦¦¦P¦¦¦20031 110145645¦20031110150803\20031110150802¦PHOENIX^^^1^C5
R¦5¦^^^AST_MIC^424940000029^CC¦INST_IN_ATTN_COMPLETE^>8¦¦¦¦¦P¦¦¦2003111 0145645¦20031110150803\20031110150802¦PHOENIX^^^1^C5
L¦1¦N
Query Upload Examples
Example of an Unsolitcited Query uploaded due to a missing organism id H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20031110151943 Q¦1¦^InvOrg^429530000072
L¦1¦N
Order Download Examples
Example of a download order for a critical panel
H¦\^&¦¦¦ P¦1
O¦1¦Critical^1¦¦^^^GNID 952¦A L¦1¦N
Query Download Examples
For test start times between Nov 10,2003 at 09:00:00 and Nov 10,2003 at 16:00:00
H¦\^& Q¦1¦¦¦¦S¦20031110090000¦20031110160000
L¦1¦N
For test result times after Nov 10,2003 at 09:00:00
H¦\^& Q¦1¦¦¦¦S¦20031110090000¦¦ L¦1¦N
For panels that have partial results and have been modded since the last request
H¦\^& Q¦1¦¦¦^INST_PARTIAL_COMPLETE^^^MOD
L¦1¦N
For the panel with 429530000072 sequence number
H¦\^&
Q¦1¦^^429530000072¦¦ L¦1¦N
For ALL accession numbers
H¦\^& Q¦1¦¦¦ALL¦¦¦
L¦1¦N
14.4. BACTEC FX
14.4.1. BACTEC FX LIS Configuration Screen
The BACTEC FX LIS Configuration Screen allows the user to specify the configurable behavior of the LIS interface. To navigate to the BACTEC FX LIS Configuration Screen select the Configuration Tab then the LIS Tab. The configurable options on the BACTEC FX LIS Configuration Screen are somewhat different than the parameters specified within the BACTEC 9000 Configuration File but are intended to give the user most of the same LIS interface functionality.
The BACTEC FX LIS interface is configured by enabling and defining fields within the LIS tab on the BACTEC FX Configuration screen. If parameters are changed on the LIS configuration screen, the new configuration will take effect immediately after the SAVE button is pressed. The BACTEC FX doesn’t need to and will not restart after the LIS configuration parameters are changed (BACTEC 9000 requires reboot after configuration changes). However BD recommends that the host LIS be inactive during the reconfiguration period. An LIS download transaction could be lost during the reconfiguration of the BACTEC FX LIS interface. In addition the BACTEC FX’s response to an active LIS query will not be transmitted or cancelled.
14.4.2. BACTEC FX Configurable Options
14.4.2.1.LIS ENABLE Group
The options for this group are mutually exclusive.
Possible parameters: ENABLED
DISABLED
Default: DISABLED
14.4.2.2.Port Parameters Group
Parameters defined in this group control the communication parameters assign to the serial communications port on the BACTEC FX that is connected to the LIS
14.4.2.2.1.BAUD
The BAUD parameter allows the adjustment of the speed of the communications link between the LIS system and the BACTEC FX.
Possible parameters:
1200
2400
4800
9600
19.2K
Default: 9600
14.4.2.2.2.PARITY
The PARITY parameter allows for the adjustment of the error check that can be performed on every character that is received and transmitted.
Possible parameters:
NONE ODD
EVEN
Default: NONE
14.4.2.2.3.DataBits
The Databits parameter allow for the adjustment of the size of the character that is transmitted and received by the Physical Protocol Layer.
Possible parameters: 7 DataBits
8 DataBits
Default: 8 DataBits
14.4.2.2.4.STOPBits
The STOPS parameter allow for the adjustment of the number of stop bits that are appended to the characters that are transmitted by the Physical Protocol Layer.
Possible parameters 1 StopBit
2 StopBits
Default: 1 StopBit
14.4.2.3.Physical Layer Group
This group allows the user to select between the physical communications protocols thst the LIS interface will use.
The options for this group are mutually exclusive.
Possible parameters: ASTM 1381
BDMODEM
Default: BDMODEM
NOTE: If the ASTM physical layer is to be used for LIS communication, refer to the BD LIS interface specification for ASTM interfaces, document number L–005933.
14.4.2.4.LIS Options Group
This group allows the user to define the LIS/LAB specific workflow(dictating when data is sent to the LIS) and the types of data that will be sent to the LIS.
14.4.2.4.1.Upload Pos Result
When this checkbox is enabled (checked), Positive results will be uploaded to the LIS, whenever a Results Upload to the LIS occurs (see LIS Solicit Results section 6.3.1.4.2)
Default: Disabled
14.4.2.4.2.LIS Solicited Results
If LIS Solicit Results is enabled (checked) Negative Results will be sent only when the LIS requests results. If Upload Positive Results is enabled Positive results will also be sent to the LIS when solicted.
If LIS Solicit Results is disabled (unchecked) Negative results will be sent automatically when the vial goes out of protocol. If Upload Positive Results is enabled Positive results will also be sent to the LIS when the vial turns Positive. If Vial Tracking is enabled then Vial Tracking information will also be sent automatically to the LIS when the vial’s location changes.
Default: Solicit
14.4.2.4.3.Vial Tracking
When this function is enabled (checked) Vial Tracking information will be sent to the LIS when the vial location changes (INST_NEW, INST_REMOVED, INST_REENTERED). This function is disabled when the LIS solicited Results checkbox is enabled (checked) so therefore, Vial Tracking information will only be sent unsolicited. The BACTEC FX will report INST_REENTERED for all vials that have been removed and reentered. The workflow for BACTEC FX allows the operator to remove an ongoing vial and reenter it.
Default: Disabled
14.4.2.4.4.Host Query Mode
The BACTEC FX will perform a Host Query to request patient demographics information from the LIS for any vial without a patient name/id that has an associated accession number. This option only works if Accession Barcoding is enabled on the Configuration screen and the LIS Solicited Results (is not selected) is operating in the Unsolicited mode.
Possible parameters:
MANUAL
User initiated Batch Host Query. When this option is selected the user can initiate a Batch Host Query via the Maintenance\Test tab by pressing the LIS button. A Batch Host Query consists of multiple individual SINGLE Host Queries for all vials in the BACTEC FX database without Patient Demographics. When this option is selected the parameter Query Timeout can also be adjusted within the LIS Options Group. The Query Timeout dictates how long the BACTEC FX will wait for the LIS to respond to a Host Query before sending the next Host Query requesting Patient data for the next vial.
SINGLE
A Single Host Query is initiated when a new vial with an accession number is entered. When this option is selected the parameter Query Timeout can also be adjusted within the LIS Options Group. The Query Timeout dictates how long the BACTEC FX will wait for the LIS to respond to a Host Query before sending the next Single Host Query requesting Patient data for the next newly entered vial.
AUTO
In Auto Mode the BACTEC FX will perform a Single Host Query when a new vial with an accession number is entered. It will also initiate a Batch Host Query based on the frequency defined in the Auto HQ Interval field. So if the LIS doesn’t respond to the Single Host Query when a vial is loaded, the BACTEC FX will automatically gather the missing patient information via a Batch Host Query sent at the next Auto HQ Interval. In AUTO mode the parameter Query Timeout can also be adjusted within the LIS Options Group. The Query Timeout dictates how long the BACTEC FX will wait for the LIS to respond to a Single Host Query before sending the next Single Host Query requesting the demographic for the next vial.
DISABLED
Do not initiate a Host Query of any kind due to instrument workflow.
Default: AUTO
Query Timout
Query Timeout is enabled for AUTO, SINGLE and MANUAL Host Query modes. It can be set from 1 – 99 seconds. The Query Timeout dictates how long the BACTEC FX will wait for the LIS to respond to a Single Host Query before sending the next Single Host Query requesting the demographic associated with the next vial.
14.4.2.4.5.Auto HQ Interval
Auto HQ Interval can be set for AUTO mode. It can be set from 1 – 999 minutes. This parameter dictates the frequency the BACTEC FX will perform an AUTO Batch Host Query.
14.4.2.4.6.Log Comms Checkbox
The Log Comms checkbox is used as a tool for BD personnel in order to assist in troubleshooting LIS communication issues. When enabled and a USB key is present, additional low-level messages are logged to a file on the USB key.
These messages, when used in tandem with the LIS messages from the event log, will assist in deciphering any communications issues.
14.4.2.4.7.Force Upload Checkbox
The Force Upload checkbox is to be used as a troubleshooting mechanism for LIS communication issues. When checked (and saved via the <Save> button on the LIS configuration tab) the Save button on the Culture Screen Vial Tab shall be disabled and the ‘Send’ button shall be displayed and enabled in its place when non-modified vial data is recalled into the Vial Tab. Pressing the ‘Send’ button will transmit a Result record of the recalled vial on the Culture screen to the LIS
14.4.2.5.ASTM 1381 Group
Not applicable for the BDModem Protocol. This group allows the user to specify Packed Frames or Unpacked Frames
Packed Frames ON/OFF
Default: ON
14.4.2.6.BD Modem Group
The parameters defined within this group are used to determine where specific data is located in the LIS records. and to define the expected characters in a message to determine specific information.
14.4.2.6.1.New Sequence Position
To ensure backwards-compatibility with LIS interfaces written prior to this specification, this option indicates where barcodes would be found in an Order record, regardless if they are downloaded or not. The NEW parameter should only be used if an LIS interface is being written or updated to this specification.
Possible parameters:
OLD NEW
Default: NEW
14.4.2.6.2.Hospital Service Field (33)
This parameter allows the selection of the field for placement of the Hospital Service Code in the Patient Record, for both download and upload. This parameter allows vendors who have implemented the old BD DMC LIS interface to use code that they already have. The DMC LIS (OLD interface) interface calls for the Hospital Service Code to be placed in field 26 of the record. The BACTEC 9000 LIS (NEW interface) interface calls for the Hospital Service Code to be placed in field 33 of the record. This option should be set to match the location that the LIS vendor will be using to provide the Hospital Service Code. If the dialog box is checked the Hospital Service Code is in field
33. If it is not checked the Hospital Service Code will be in field 26 Possible parameters: 26 or 33
Default: 33
14.4.2.6.3.SOH
The SOH parameter allows the selection of the character that will be used to indicate the Start Of Header for a message. The values are specified in hex.
Possible parameters: 00hex..FFhex Default: 01hex
14.4.2.6.4.EOT
The EOT parameter allows the selection of the character that will be used to indicate the End Of Text for a message. The values are specified in hex. The value that is used by the LIS vendors could be one of two possible values. The protocol has been implemented with EOT set to 2 or 4. When the LIS system is configured the LIS vendor should indicate the value that they are using for EOT. The EOT value in the configuration file should be set accordingly.
Possible parameters: 00hex..FFhex Default: 04hex
14.4.2.6.5.ACK
The ACK parameter allows the selection of the character that will be used to indicate the Acknowledgement for a message. The values are specified in hex.
Possible parameters: 00hex..FFhex Default: 06hex
14.4.2.6.6.NAK
The NAK parameter allows the selection of the character that will be used to indicate the Negative Acknowledgement for a message. The values are specified in hex.
Possible parameters: 00hex..FFhex Default: 15hex
14.4.2.6.7.CAN
The CAN parameter allows the selection of the character that will be used to indicate the Cancellation for a message. The values are specified in hex.
Possible parameters: 00hex..FFhex Default: 18hex
14.4.2.6.8.SYN
The SYN parameter allows the selection of the character that will be used to indicate the Synchronization request for the start of a message. This control character has been implemented using two different values. The LIS vendor should be queried as to the value they used in the implementation and the option set accordingly. The two values that have been used are 26 and 16 hex. The values are specified in hex.
Possible parameters: 00hex..FFhex Default: 16hex
14.4.2.7.Send Button
When vial information for a non-pending, non-orphan sequenced vial is recalled onto the Culture screen, the <Send> button shall become enabled – when the Force Upload Checkbox is enabled (see section 14.4.2.4.7).
Touching the <Send> button shall cause the displayed vial data to be transmitted as an LIS message immediately. Since this is used as a troubleshooting mechanism, the information will be formulated into the appropriate LIS message format and the LIS message to be transferred to the attached LIS.
Once any data on the Culture Screen Vial Tab has been modified, the ‘Send’ button shall be removed from the display and the ‘Save’ button shall be displayed and disabled.

## Table 23
14.4.2.8.Print Button
The Print button on the Maintenance/Configuration/Lis tab shall be displayed and enabled when the LIS interface is enabled.
When enabled, the Print button shall be accessible by all users; no password or enabling of the Log Comm checkbox is required in order for this button to be operational.
Touching the Print button shall cause all of the LIS messages that are contained in the event log to be printed to the selected printer in the selected language. This will include the messages time stamp in order for the timing of these messages as well as the low-level messages being saved on the USB key to be coordinated.
14.4.3. Message Content
14.4.3.1.Field List
The BACTEC FX shall exchange fields with the LIS per the following table:
Header Record Field Name | ASTM Pos. | Direction
Sender Name | H, 5, 1, 1 | U
Version Number | H, 13, 1, 1 | U
Message Date/Time | H 14, 1, 1 | U
Patient Record Field Name | ASTM Pos. | Direction
Patient ID | P, 4, 1, 1 | U/D
Patient Last Name | P, 6, 1, 1 | U/D
Patient First Name | P, 6, 1, 2 | D
Patient Middle Name | P, 6. 1, 3 | D
Patient Name Suffix | P, 6, 1, 4 | D
Patient Name Title | P, 6, 1, 5 | D
Old Hospital Service | P, 26, 1, 1 | U/D
Hospital Service | P, 33, 1, 1 | U/D

## Table 24
Order Record Field Name | ASTM Pos. | Direction
Accession Number | O, 3, 1, 1 | U/D
Old Sequence 1 | O, 3, 1, 3 | D
Old Sequence 2 | O, 3, 2, 3 | D
Old Sequence 3 | O, 3, 3, 3 | D
Old Sequence 4 | O, 3, 4, 3 | D
Old Sequence 5 | O, 3, 5, 3 | D
Old Sequence 6 | O, 3, 6, 3 | D
Old Sequence 7 | O, 3, 7, 3 | D
Old Sequence 8 | O, 3, 8, 3 | D
Old Sequence 9 | O, 3, 9, 3 | D
Old Sequence 10 | O, 3, 10, 3 | D
Test ID | O, 5, 1, 4 | U
Sequence 1 | O, 5, 1, 5 | D
Sequence 2 | O, 5, 2, 5 | D
Sequence 3 | O, 5, 3, 5 | D
Sequence 4 | O, 5, 4, 5 | D
Sequence 5 | O, 5, 5, 5 | D
Sequence 6 | O, 5, 6, 5 | D
Sequence 7 | O, 5, 7, 5 | D
Sequence 8 | O, 5, 8, 5 | D
Sequence 9 | O, 5, 9, 5 | D
Sequence 10 | O, 5, 10, 5 | D
Collection Date Time | O, 8, 1, 1 | U/D
Result Record Field Name | ASTM Pos. | Direction
Result ID Code | R, 3, 1, 4 | U
Result Sequence Number | R, 3, 1, 5 | U
Test Status | R, 4, 1, 1 | U
Prelim Status | R, 9, 1, 1 | U
Start Date Time | R, 12, 1, 1 | U
End Date Time | R, 13, 1, 1 | U
Instrument Type | R, 14, 1, 1 | U
Media Type | R, 14, 1, 2 | U
Protocol Length | R, 14, 1, 3 | U
Instrument Number | R, 14, 1, 4 | U
Instrument Location | R, 14, 1, 5 | U
Request Record Field Name | ASTM Pos. | Direction
Request Starting Accession Number | Q, 3, 1, 2 | U
Request Test Status | Q, 5, 1, 2 | D
Starting Date/Time | Q, 7, 1, 1 | D
Ending Date/Time | Q, 8, 1, 1 | D
Request Information Status Code | Q, 13, 1, 1 | U/D
Terminator Record Field Name | ASTM Pos. | Direction
Termination Code | L, 3, 1, 1 | U/D
14.4.3.2.Field Descriptions
14.4.3.2.1.Header Fields
General – Same as EpiCenter (see section 8.2.1)
14.4.3.2.2.Patient Fields
Patient ID (P, 4, 1, 1) - Identifier that uniquely identifies a patient. This is a required field for patient demographic downloads. This field can be up to 16 characters long. (e.g. PatId123)
Patient Name (P, 6, 1, 1-5) – Patient name is divided into 5 components in the ASTM E_1394 specification. The BACTEC FX stores these fields as a single string in its database. Therefore the BD LIS Interface concatenates these fields from a download message. The patient name can be up to 40 characters long. If these fields are selected for upload on the BACTEC FX the whole name is placed in the first component of this field (P, 6, 1, 1). Patient name shown in the sample record is Dr. John R Doe Jr.
Patient Last Name (P, 6, 1, 1)
Patient First Name (P, 6, 1, 2)
Patient Middle Name (P, 6, 1, 3)
Patient Name Suffix (P, 6, 1, 4)
Patient Name Title (P, 6, 1, 5)
Old Hospital Service (P, 26, 1, 1) – (Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) An acronym or abbreviation for the ward or hospital department involved in caring for this patient. (e.g. ER may represent the Emergency Room). (Old Hospital Service field only for compatibility with older BACTEC 9000 LIS interfaces).
Hospital Service (P, 33, 1, 1) – An acronym or abbreviation for the ward or hospital department involved in caring for this patient. (e.g. ER may represent the Emergency Room).
Accession Number (O, 3, 1, 1) - Maximum field length: 20 characters. A unique alphanumeric string that identifies a specimen.
Old Sequence1 (O, 3, 1, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence2 (O, 3, 2, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence3 (O, 3, 3, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence4 (O, 3, 4, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable..
Old Sequence5 (O, 3, 5, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence6 (O, 3, 6, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence7 (O, 3, 7, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence8 (O, 3, 8, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence9 (O, 3, 9, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Old Sequence10 (O, 3, 10, 3) –(Do Not Use unless the LIS vendor has previously implemented the old BD DMC LIS interface and that code is reused) 12 digit numeric string identifying the instrument consumable.
Test Id (O, 5, 1, 4) –BACTEC FX transmits “BACTECFX_GND” for the Test ID.
Sequence1 (O, 5, 1, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence2 (O, 5, 2, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence3 (O, 5, 3, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence4 (O, 5, 4, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence5 (O, 5, 5, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence6 (O, 5, 6, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence7 (O, 5, 7, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence8 (O, 5, 8, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence9 (O, 5, 9, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Sequence10 (O, 5, 10, 5) -- 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Collection Date Time (O, 8, 1, 1) Date and time that the specimen was collected from the patient, formatted as described in ASTM E_1394 section 6.6.2.
14.4.3.2.3.Result Fields
Result ID Code (R, 3, 1, 4) The BACTEC FX transmits a “GND” as the Result ID Code.
Result Sequence Number (R, 3, 1, 5) – 12 digit numeric string identifying the instrument (BACTEC FX) consumable.
Test Status LIS Code (R, 4, 1, 1) -- This describes the Status of a BACTEC FX vial. The possible values are:
 INST_NEW
 INST_POSITIVE
 INST_NEGATIVE
 INST_REMOVED
 INST_REENTERED
 INST_ONGOING
Preliminary Status (R, 9, 1, 1) -- We always put a P in this field, because all results from BACTEC FX are preliminary.
Start Date/Time (R, 12, 1, 1) This is the date and time that the vial was physically placed in the instrument. This field is represented in the YYYYMMDDHHMMSS format.
End Date/Time (R, 13, 1, 1) This is the date and time that the panel’s status transitioned to POSITIVE or OUT OF PROTOCOL. This field is represented in the YYYYMMDDHHMMSS format.
Instrument Type (R, 14, 1, 1) The BACTEC FX transmits BACTECFX for the instrument Type.
Media Type(R, 14, 1, 2) 2 digit code representing the type of media contained within a BACTEC vial.
Protocol Length(R, 14, 1, 3) this field indicates the length of time required to complete the BACTEC GND test. For BACTEC FX(a growth and detection test), this number represents the number of days before a vial can be considered negative.
Instrument Number (R, 14, 1, 4) Instrument number between 1 and 99.
Instrument Location (R, 14, 1, 5) Indicates the vial position inside the instrument as an 8 character string = ii-d-rcc
where:
ii = instrument number(1-99) d = drawer letter(A-D)
r = row letter(A-K with no I)
cc= column number (1-10)
14.4.3.2.4.Request Fields
The ASTM 1394 protocol defines this as a Request Information Record message. A download Request record is also known as a “Query”. The only type of download “Query” request that the BACTEC FX instrument accepts is a request for test results. A Terminator record with a termination code of “F”(processed) or “Q”(error) is used to signal the last packet of a query response. If a query request was invalid, the instrument interface responds with only the Terminator record with a termination code of “Q”. If the instrument interface cannot locate any specimens in the active or history databases that meet the query criteria, it responds with only the Terminator record with a termination code of “F”.
A query contains a set of request parameters that are used to determine which specimen(s) to access. The following tables define the ASTM 1394 Request record fields that can be transmitted from the LIS in a Download message and are accepted and acted upon by the BACTEC FX.
Test Status (Q, 5, 1, 2) The BACTEC FX instrument accepts queries for vials with a status as follows:
INST_POSITIVE = POSITIVE & Manual Positive vials INST_NEGATIVE = NEGATIVE & Manual Negative vials
ALL = POSTIVE & NEGATIVE & ONGOING & Manuals Positive & Manual Negative vials
AWL = (Do Not Use unless the LIS vendor has previously implemented the BD DMC LIS interface and that code is reused) POSTIVE & NEGATIVE & ONGOING & Manuals Positive & Manual Negative vials
Starting Date\Time (Q, 7, 1, 1) – Requested starting time for query -- the beginning of the Date/Time range to gather data for. If it is null the BACTEC FX will use the Date 1/1/1970.
End Date\Time (Q, 8, 1, 1) – Requested Ending time for the query -- the ending of the Date \Time range to gather data for. It must be greater than the Starting Date\Time or the query will be rejected. If it is null, the BACTEC FX will use the current date time.
When a "Time Range" has been specified in a Query (Begin Time and/or End Time fields), the time field from the instrument's database that is used for comparison is dependent upon the Query's "Test Status" field as follows,
INST_POSITIVE -- Time the specimen went positive. INST_NEGATIVE -- Time the specimen went out of protocol.
INST_ONGOING – The time the specimen was entered into the BACTEC FX. MANUAL vials – Time the user set the Status.
14.4.3.2.5.Terminator Fields
General – Same as EpiCenter (see section 8.2.7)
14.4.4. Example ASTM 1394 Protocol
The following are examples of various BACTEC FX message transmissions that conform to the Becton Dickinson Common LIS Interface using the ASTM 1394 protocol.
BACTEC FX G&D specimen that has been entered into an instrument as a new vial: (note: whether the vial has received a first test or not is not relevant in the BACTEC FX instrument.) – sent from the BACTEC FX instrument to the LIS.
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20081017105509 P¦1
O¦1¦ACC-001¦¦^^^BACTECFX_GND¦¦¦19700101000000 R¦1¦^^^GND^446500000000¦INST_NEW¦¦¦¦¦P¦¦¦20081017105508¦¦BACTECFX^65^5^01^A- G01
L¦1¦N
BACTEC FX G&D specimen that has been tested POSITIVE – sent from the BACTEC FX instrument to the LIS:
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20081024064721 P¦1¦¦PATID¦¦LastName FirstName
O¦1¦ACC-001¦¦^^^BACTECFX_GND¦¦¦19700101000000 R¦1¦^^^GND^446500000000¦INST_POSITIVE¦¦¦¦¦P¦¦¦20081017105508¦20081017110031¦BA CTECFX^65^5^01^A-G01
L¦1¦N
BACTEC FX G&D specimen that has been tested NEGATIVE – sent from the BACTEC FX instrument to the LIS:
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦¦V1.0¦20081017110336 P¦1¦¦PATID2¦¦LastName2 FirstName2 O¦1¦DAG¦¦^^^BACTECFX_GND¦¦¦19700101000000
R¦1¦^^^GND^446504178134¦INST_NEGATIVE¦¦¦¦¦P¦¦¦20080918155201¦20080923160143¦BA CTECFX^65^28^01^A-A02
L¦1¦N
A Request Record (Query) sent from an LIS to a BACTEC FX instrument requesting all specimen results (all positive, negative, and ongoing vials records will be sent).
H¦\^&
C¦1¦This is a request for POSITIVE, NEGATIVE AND ONGOING vials from the LIS Q¦1¦ALL^ALL¦¦^ALL¦¦¦¦¦¦¦¦P
L¦1
A Request Record (Query) sent from an LIS to a BACTEC FX instrument requesting only POSITIVE specimen results for a period of time (vials that had a positive result called between 9/23/2008 4:01:43 p.m. and 10/17/2008 10:55:08 a.m.)
H¦\^&
C¦1¦This is a request for POSITIVE vials from the LIS Q¦1¦ALL^ALL¦¦^INST_POSITIVE¦¦20080923160143¦20081017105508¦¦¦¦¦P
L¦1
A Patient Record sent from an LIS to a BACTEC FX instrument
H¦\^&
C¦1¦L¦Patient information from LIS to BACTEC FX P¦1¦¦PATID3¦¦LastName^FirstName^Middle^Suffix^Title¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦ HosSer¦
O¦1¦Acc-003¦ ¦^^^\^^^\^^^\^^^\^^^\^^^\^^^\^^^\^^^\^^^\^^^¦¦¦20081025120000 L¦1
14.5. Innova
14.5.1. LIS Connection
The instrument attempts to connect to an LIS server when processing of a batch begins. If a connection to the LIS server has not yet been established or has been disconnected, the instrument will attempt to establish a connection when processing of subsequent batches begin.
The instrument connects to the LIS via Ethernet TCP/IP using a configurable IP address and port. Configuring the IP address and port is done via a CSV file.
The instrument exchanges data with the LIS per the protocol specified in the Data Link Section of ASTM E_1381.
The instrument formats the content of the data exchanged as messages containing a collection of records per ASTM E_1394.
The instrument will send and receive frames that contain only a single logical record (i.e. Header, Patient, Order, Result and Terminator records all are sent in separate frames).
The instrument will not respond to queries generated by the LIS. The instrument will not accept unsolicited data sent from the LIS.
The implemented LIS Interface specification is selectable to be either the legacy interface or the preferred BD ASTM interface. The interface is configurable via a CSV file.
Selecting the BD ASTM Interface will result in the following changes:
 Queries to the LIS will not contain a confirmation record.
 Specimen plating status will be sent separately formatted as an H/P/O/R/L message.
The instrument will expect a query response from the LIS within a configurable amount of time. The timeout value is configurable via a CSV file. The default time is 15 seconds.
14.5.2. BD ASTM Interface
14.5.2.1. Overview
The instrument will default to the BD ASTM Interface.
When the instrument reads a new container ID barcode from a container, a query for patient demographics and test orders is sent to the LIS.
When processing has successfully completed for a container, the instrument will send a result message to the LIS containing H/P/O/R/L records.
Sending of the result message can be enabled or disabled via a CSV file. The default setting is enabled.
14.5.2.2. Valid Characters
The following characters are valid at the Innova instrument.
A – Z, a – z, 0 – 9,; : , . / @ $ % ( ) _ + - ` =
14.5.2.3. Invalid Characters
The following characters are invalid at the Innova instrument.
" * ? [ ] ! # ¦ ' & < > { } ~ ^ \
These characters cannot be entered at the Innova instrument but could be supplied by and LIS/EpiCenter Server download.
If these characters are encountered in the Patient Information record that has been received from the LIS/EpiCenter server, the following will occur :
 If an illegal character is contained in the container barcode field (O, 3, 1, 1), the illegal character will be replaced with a “$”. The barcode will be passed on for further validation. When the “Mismatched barcode” check is performed, the check will fail and the container will be rejected.
If an illegal character is contained in any other field of the record, the field will be “nulled out.” This will result in several potential scenarios:
 If the Innova is operating in LIS/EpiCenter Protocol Defined Source mode and if the Specimen Type/Body Site field has been nulled out, the specimen will be rejected with a protocol not found.
 If the field that has been nulled out is contained in the label configuration, this field will be null on the plate label.
 If the field is not printed on the plate label and it is not the barcode field or the Specimen Type/Body Site field in LIS/EpiCenter Protocol Defined Source mode, then the user will never see this nulled - out field.
 The result record that is echoed back to the LIS will contain the nulled - out field. This should cause no issues in either the LIS or the EpiCenter Server sites.

## Table 25
14.5.2.4. Message Content
The following sections describe which fields are exchanged with the Innova instrument.
14.5.2.5. Field List
Each field is displayed with its default mapping in the ASTM E_1394 records. This position consists of a record type (Patient, Order, Result, Query, or Comment), a field delimiter counter, repeat delimiter counter, and component delimiter counter. The fields are grouped according to the ASTM E_1394 record type they are mapped to.
The field list also indicates if that field is defaulted as an upload (U) field, download (D) field, or both (U/D).
The first two fields of every record include the Record Type indicator and a record index value. These fields are considered part of every record but are not listed in these tables.
Header Record Field Name | ASTM Pos. | Direction
Delimiter Field | H, 2, 1, 1 | U/D
Sender Name | H, 5, 1, 1 | U/D
Processing ID | H, 12, 1, 1 | U/D
Version Number | H, 13, 1, 1 | U/D
Message Date/Time | H 14, 1, 1 | U/D
Patient Record Field Name | ASTM Pos. | Direction
Patient ID | P, 4, 1, 1 | U/D
Patient Name (Last) | P, 6, 1, 1 | U/D
Patient Name (First) | P, 6, 1, 2 | U/D
Patient Name (Middle) | P, 6, 1, 3 | U/D
Date of Birth | P, 8, 1, 1 | U/D
Patient Sex | P, 9, 1, 1 | U/D
Hospital Service | P, 33, 1, 1 | U/D
Hospital Client | P, 34, 1, 1 | U/D
Order Record Field Name | ASTM Pos. | Direction
Accession Number / Container ID* | O, 3, 1, 1 | U/D
Culture ID | O, 3, 1, 2 | U/D
Barcode Number / Accession Number* | O, 5, 1, 5 | U/D
Collection Date/Time | O, 8, 1, 1 | U/D
Specimen Type | O, 16, 1, 1 | U/D
Body Site | O, 16, 1, 2 | U/D
Ordering Physician | O, 17, 1, 1 | U/D
Server Free Text | O, 19, 5, 1 | U/D
Report Type | O, 26, 1, 1 | U/D

## Table 26
*See Field Descriptions - Order Fields Subsection
Result Record Field Name | ASTM Pos. | Direction
Result ID Code | R, 3, 1, 4 | U
Test Status | R, 4, 1, 1 | U
Mode | R, 4, 1, 9 | U
Preliminary/Final Status | R, 9, 1, 1 | U
Media Setup Date/Time | R, 13, 1, 2 | U
Instrument Type | R, 14, 1, 1 | U
Machine Instrument Number | R, 14, 1, 4 | U
Instrument Location | R, 14, 1, 5 | U
Protocol Name | R, 14, 1, 8 | U
Request Record Field Name | ASTM Pos. | Direction
Accession Number | Q, 3, 1, 2 | U
Container Barcode Number | Q, 3, 1, 3 | U
Request Instrument Number | Q, 5, 1, 4 | U
Request Information Status Code | Q, 13, 1, 1 | U
Comment Record Field Name | ASTM Pos. | Direction
Comment Text | C, 4, 1, 1 | U
Comment Type | C, 5, 1, 1 | U
Terminator Record Field Name | ASTM Pos. | Direction
Termination Code | L, 3, 1, 1 | U/D
14.5.2.6. Field Descriptions
This section provides detailed information for each field listed in the previous section.
14.5.2.6.1. Header Fields
The header record that starts each logical message is the same from all BD instruments. An example record is provided below.
H¦\^&¦ ¦ ¦Becton Dickinson¦ ¦ ¦ ¦ ¦ ¦ ¦P ¦X1.00¦20100901184200
Delimiter Fields (H, 2, 1, 1) – These single characters can be used to process the remainder of the message. These characters denote the field, repeat, component and escape delimiters as described in the ASTM E_1394 specification.
These fields are defaulted to “¦”, “\”, “^”, and “&” respectively, as shown above.
Sender Name (H, 5, 1, 1) – Messages coming from a Becton Dickinson instrument have the BD identifier in this field as shown above.
Processing ID (H, 12, 1, 1) – Indicates how this message is to be processed. The following codes will be used:
P – Treat message as an active message to be completed according to standard processing.
Version Number (H, 13, 1, 1) – This version number represents the version of the LIS interface used for communications. For the Innova instrument, the version number will begin with X1.00.
Message Date/Time (H, 14, 1, 1) – BD includes the current time and date when constructing his message, formatted as described in the ASTM E_1394 specification in section 6.6.2.
14.5.2.6.2. Patient Fields
P ¦ 1 ¦ ¦ PatId123 ¦ ¦ Last Name ^ First Name^ MI¦ ¦19651029
¦M¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦Hospital Service¦Hospital Client
Patient ID (P, 4, 1, 1) - Identifier that uniquely identifies a patient. This field can be up to 16 characters long (e.g. PatId123). This field is optional.
Patient Name (P, 6, 1, 1-3) – Patient name is divided into 5 components in the ASTM E_1394 specification. The Innova instrument will accept the first 3 components of this field. The instrument stores these fields as a single string in its database. The patient name can be up to 40 characters long. When uploaded, the whole name is placed in the first component of this field (P, 6, 1, 1). This field is optional.
Date of Birth (P, 8, 1, 1) – Patient date of birth formatted as described in ASTM E_1394 section 6.6.2. The date shown in the sample record is Oct. 29, 1965. This field is optional.
Patient Sex (P, 9, 1, 1) – There are 3 predefined values accepted for this (M)ale, (F)emale, and (U)nspecified. If the value does not match the code for either Male or Female then the sex is set to Unspecified. This field is optional.
Hospital Service (P, 33, 1, 1) – This is the location within the hospital that the specimen came from, for instance ER (emergency room), OR1 (operating room 1). This field is optional.
Hospital Client (P, 34, 1, 1) – This is the facility that the specimen came from, for instance Smith Radiology or Central Labs. This field is optional.
14.5.2.6.3. Order Fields
O ¦ 1 ¦ Access1^331316^ACC123¦ ¦ ^^^^ ¦ ¦ ¦ 20090108080510 ¦ ¦ ¦ ¦
¦ ¦ ¦ ¦ URINE^CleanCatch ¦ Dr. Ameigh ¦ ¦ \ \ \ \Server Free Text
¦ ¦ ¦ ¦ ¦ ¦ ¦ O
Accession Number (O, 3, 1, 1) – A unique alphanumeric string that identifies a specimen. This field can be up to 16 characters long. This field is required.
NOTE: if “CONTAINER_BC_INSTEAD_OF_ACC” flag is set to 1 (one – CID Mode 1), this field will contain the container barcode that was queried for to obtain the accession number.
Culture ID (O, 3, 1, 2) – An alphanumeric string that identifies a culture. This field is optional.
Barcode Number (O, 5, 1, 5) – An alphanumeric string to be included on the printed plate label. This field is optional.
NOTE: if “CONTAINER_BC_INSTEAD_OF_ACC” flag is set to 1 (one – CID Mode 1), this field will contain the accession that is to be associated to the CID that was queried for in Q, 3, 1, 3 field.
For additional information on “CONTAINER_BC_INSTEAD_OF_ACC” functionality, see the CONTAINER_BC_INSTEAD_OF_ACC section.
Collection Date/Time (O, 8, 1, 1) – Date and time that the specimen was collected from the patient, formatted as described in ASTM E_1394 section
6.6.2. This field is optional.
Specimen Type (O, 16, 1, 1) – The type of specimen collected from the patient. This field is optional.
For LIS/EpiCenter Defined Protocol Source, this field must exactly match the field at the Innova associated to a protocol. This field is a maximum of 14 characters at the Innova.
Body Site (O, 16, 1, 2) – The body site for the source of the specimen collected from the patient. This field is optional.
For LIS/EpiCenter Defined Protocol Source, this field must exactly match the field at the Innova associated to a protocol. This field is a maximum of 14 characters at the Innova.
Ordering Physician (O, 17, 1, 1) – The physician that has ordered the test. This field is optional. This field has a maximum of 16 characters when it is printed on the label at the Innova.
Server Free Text (O, 19, 5, 1) – This field contains free text that can hold any alphanumeric information (such as special instructions to look for fungi on the plate, etc.). This field will have a maximum of 16 characters when printed on the label. This field is optional.
Report Types (O, 26, 1, 1) – If the Report Types field contains an “X”, “Y”, or “Z”, the specimen is rejected. Any other value will be ignored. This field is optional. (From ASTM 1384 specifications)
X – request cancelled
Y – no order on record for this test (in response to query)
Z – no record of this patient (in response to query)
14.5.2.6.4. Result Fields
R ¦ 1 ¦ ^^^STREAK¦INST_COMPLETE^^^^^^^^MANAGED ¦ ¦ ¦ ¦
¦ P ¦ ¦ ¦ ¦^20110817130534 ¦ INNOVA^^^1^1-1-1^^^Urine
Result ID Code (R, 3, 1, 4) – This is a BD defined code that indicates the type of information being exchanged in the result record. This field contains the value, “STREAK”, to indicate the results are from an automated plate streaking instrument.
Test Status (R, 4, 1, 1) – The status value for a particular specimen. This field is the pre-defined value INST_COMPLETE for specimens that have completed successfully. For specimens that have not completed successfully, the value is INST_ERROR.
Mode (R, 4, 1, 9) – The mode indicates whether the specimen was run with the “User Defined Protocol” mode of operation (UD) or the LIS Defined Protocol mode of operation (LD).
This field will contain the following values: “MANAGED” – specimen processed in LD mode. “BATCH” - specimen processed in UD mode
Preliminary/Final Status (R, 9, 1, 1) – This field contains a “P” to indicate the results should be considered Preliminary Results.
Media Setup Date/Time (R, 13, 1, 2) – This field indicates the time the specimen was successfully streaked. If the specimen was not successfully streaked, this field will be blank.
Instrument Type (R, 14, 1, 1) – This field indicates which BD instrument produced the result. This field contains the value, “INNOVA”.
Machine Instrument Number (R, 14, 1, 4) – This is the number assigned to the instrument.
Instrument Location (R, 14, 1, 5) – Indicates the location of the container in the instrument. Represented as drawer number - canoe number - position in canoe.
Protocol Name (R, 14, 1, 8) – This is the protocol used to plate the specimen. It is derived from the Specimen Type/Body Site supplied by the LIS or selected by the user for a User Defined specimen.
14.5.2.6.5. Request Fields
Standard Query
Q ¦ 1 ¦ ^0123456¦ ¦^^^1 ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦O¦
CID Mode 1 Query
Q ¦ 1 ¦ ^^0123456¦ ¦^^^1 ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦O¦
Accession Number (Q, 3, 1, 2) – The unique alphanumeric string that identifies a specimen. This field can be up to 16 characters long. In the example above, the Accession number is “0123456”.
Container Barcode Number (Q, 3, 1, 3) – The unique alphanumeric string that identifies a container. This field can be up to 16 characters long.
This field will ONLY be populated when the Innova has its “CONTAINER_BC_INSTEAD_OF_ACC” flag set to 1 (one – CID Mode 1).
NOTE: when this field is used, its value MUST be returned in O, 3, 1, 1 so that the Innova can verify the accession number that is returned is intended for the container barcode that was queried.
Request Information Status Code (Q, 13, 1, 1) – The following codes will be used:
O – Requesting test orders and demographics only (no results)
A – abort/cancel last request criteria (allows a new request to follow, sent from Innova if query result not received before timeout)
14.5.2.6.6. Comment Fields
The comment record in the Innova will immediately follow the result record if the result is INST_ERROR. It will contain the enum value associated to the specimen’s result and the associated string that is displayed to the user on the DB Log screen and on the Batch Results screen. If there is more than one result enum value associated to the Specimen, then more than one comment record will be sent to the LIS/EpiCenter, one for each result enum value and string.
C ¦ 1 ¦ ¦<4> No Barcode ¦ R
C ¦ 2 ¦ ¦<28> Low Liquid Level ¦ R
Comment Text (C, 4, 1, 1) – The text to be used for the patient, specimen or isolate comment. This field can be up to 1600 characters long.
The number between the < > will be the enum value that is associated to the Innova ContainerProcessingStatusMask. The text that follows in this field will be the string that is displayed to the user at the Innova and that is associated to the ContainerProcessingStatusMask enum value.
Comment Type (C, 5, 1, 1) – The type of comment record. This field should be a P, S, I, E, T, or R for patient, specimen, isolate, chartables rule, special message comments, or result comment respectively.
For the Innova, this field will contain an “R”.
14.5.2.7. Examples
14.5.2.7.1. Query - Standard
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦P¦X1.00¦20100816171740 Q¦1¦^10467¦¦^^^ALL¦¦¦¦¦¦¦¦O
L¦1¦N
14.5.2.7.2. LIS Query Response - Standard
H¦\^&¦¦¦LIS¦¦¦¦¦¦¦P¦1¦20100806093245
P¦1¦¦PatID1¦¦Clark^John¦ ¦19800806 ¦M¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦ER¦Universal Labs O¦1¦10467¦¦¦¦¦20100805093245¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦URINE^Clean Catch¦
L¦1¦F
14.5.2.7.3. Query - CID Mode 1
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦P¦X1.00¦20100816171740 Q¦1¦^^10467¦¦^^^ALL¦¦¦¦¦¦¦¦O
L¦1¦N
14.5.2.7.4. LIS Query Response – CID Mode 1
H¦\^&¦¦¦LIS¦¦¦¦¦¦¦P¦1¦20100806093245
P¦1¦¦PatID1¦¦Clark^John¦ ¦19800806 ¦M¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦ER¦Universal Labs O¦1¦10467 ¦¦^^^^PatAccession¦¦¦20100805093245¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦URINE^Clean Catch¦ L¦1¦F
14.5.2.7.5. LIS Query Response with no information
H¦\^&¦¦¦LIS¦¦¦¦¦¦¦P¦1¦20100806093245 L¦1¦I
14.5.2.7.6. Confirmation Message Upload
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦P¦X1.00¦20100816172135 P¦1¦¦PatID1¦¦John Clark¦ ¦19800806 ¦ M ¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦ER¦Universal Labs O¦1¦10467¦¦^¦¦¦20100805093245¦¦¦¦¦¦¦¦URINE ^ Clean Catch
R¦1¦^^^STREAK¦INST_COMPLETE¦ ¦ ¦ ¦ ¦P¦ ¦ ¦ ¦ ¦INNOVA^^^1^3-2- 6^^^Plate Protocol
L¦1¦N
14.5.2.7.7. Confirmation Message Upload Error Status
H¦\^&¦¦¦Becton Dickinson¦¦¦¦¦¦¦P¦X1.00¦20100816172135 P¦1¦¦PatID1¦¦John Clark¦ ¦19800806 ¦ M ¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦¦ER¦Universal Labs O¦1¦10467¦¦^¦¦¦20100805093245¦¦¦¦¦¦¦¦URINE ^ Clean Catch
R¦1¦^^^STREAK¦INST_ERROR¦ ¦ ¦ ¦ ¦P¦ ¦ ¦ ¦ ¦INNOVA^^^1^3-2-6^^^Plate Protocol
C ¦ 1 ¦ ¦<4> No Barcode ¦ R
C ¦ 2 ¦ ¦<28> Low Liquid Level ¦ R L¦1¦N
14.5.2.7.8. Host Reject Messages
14.5.2.7.8.1. Report Type X
Test was deleted do not plate.
H ¦ \ ^ & ¦ ¦ ¦ HOST ¦ ¦ ¦ ¦ ¦ ¦ ¦ P ¦ 1 ¦ 20090109100400 P ¦ 1 ¦ ¦ 13099844202 ¦ ¦^¦ ¦ ¦
O ¦ 1 ¦ 13099844202 ^ ^ ¦ ¦ ^ ^ ^ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ X
L ¦ 1¦N
14.5.2.7.8.2. Report Type Y
No specimen information found – if this is LIS Driven mode, do not plate no protocol information available.
H ¦ \ ^ & ¦ ¦ ¦ HOST ¦ ¦ ¦ ¦ ¦ ¦ ¦ P ¦ 1 ¦ 20090109100400 P ¦ 1 ¦ ¦ 13099844202 ¦ ¦^¦ ¦ ¦
O ¦ 1 ¦ 13099844202 ^ ^ ¦ ¦ ^ ^ ^ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ Y
L ¦ 1¦N
14.5.2.7.8.3. Report Type Z
No patient information found – if this is LIS Driven mode and there is no associated Specimen information, do not plate no protocol information available.
H ¦ \ ^ & ¦ ¦ ¦ HOST ¦ ¦ ¦ ¦ ¦ ¦ ¦ P ¦ 1 ¦ 20090109100400 P ¦ 1 ¦ ¦ 13099844202 ¦ ¦^¦ ¦ ¦
O ¦ 1 ¦ 13099844202 ^ ^ ¦ ¦ ^ ^ ^ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ ¦ Z
L ¦ 1¦N
14.5.2.8.CONTAINER_BC_INSTEAD_OF_ACC Barcode
This section describes the CONTAINER_BC_INSTEAD_OF_ACC barcode functionality that will be available in a Post-V3.10 release. This functionality gives the user the ability to read a barcode on the specimen container and query the LIS for the associated Accession barcode.
A new parameter will be added to the MachineGeneralAttributes.csv file. This parameter will be named CONTAINER_BC_INSTEAD_OF_ACC. Its default value will be “0” for standard container barcode query. If the barcode on the container is not the accession and the user desires to have the accession returned, this parameter should be configured to have a value of 1 (one).
14.5.2.8.1.Standard Container Barcode is Accession Functionality
If the barcode that is on the container IS the Accession barcode, the LIS/EpiCenter will be queried for Patient Information by sending the container barcode that was read in Q, 3,1, 2. When the LIS/EpiCenter returns the Patient Information, it must return the same barcode that was queried for in O,
3, 1, 1. If the same barcode is not returned in this O, 3, 1, 1; the
Patient Information will be rejected with a “Mismatched Barcode” error.
14.5.2.8.2.CONTAINER_BC_INSTEAD_OF_ACC Functionality – CID
Mode 1
If the barcode that is on the container IS the Container barcode and NOT the Accession barcode, the LIS (and in a future EpiCenter release), will be queried for Patient Information by sending the container barcode that was read in Q, 3, 1, 3. When the Patient Information is returned, the barcode that was queried for must be returned in O, 3, 1, 1 and the associated accession must be returned in O, 5, 1, 5. If the same barcode is not returned in this O, 3, 1, 1; the Patient Information will be rejected with a “Mismatched Barcode” error.